## Coriolis Flowmeter VDR
Vendor Document Requirement (VDR) — Coriolis Flowmeter — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Applicability | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes |  | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes |  | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration & witness points | Project Admin & Planning | All | Include factory mass/density calibration, zero verification, functional tests | R/W | Yes |  | If requested | Yes | Yes
4 | Manufacturing / Assembly Schedule | Project Admin & Planning | All | Milestones and submission plan | R | Yes |  | No | Yes | No
5 | Nonconformance/Concession/Deviation & Corrective Action templates | Project Admin & Planning | All |  | R | Yes |  | No | Yes | No
6 | Instrument Datasheet — Coriolis Flowmeter (per tag) | Design & Engineering | All | Model, tube geometry (U/Ω/straight), size/rating, end connection, outputs/comms, power, Ex rating | R | Yes |  | Yes | Yes | Yes
7 | Application & Sizing Check | Design & Engineering | All | Fluid, viscosity/solids %, pressure/temperature window, expected ranges (mass, density, temp), minimum flow, pressure drop | R | Yes |  | Yes | Yes | Yes
8 | Accuracy & Performance Statement — Mass Flow | Design & Engineering | All | Reference accuracy, linearity, repeatability, stability, turndown | R | Yes |  | Yes | Yes | Yes
9 | Accuracy & Performance Statement — Density | Design & Engineering | All | Reference accuracy, temp compensation, stability | R | Yes |  | Yes | Yes | Yes
10 | Zero Stability / Zero Verification Specification | Design & Engineering | All | Zero stability, procedure, allowable drift | R | Yes |  | Yes | Yes | Yes
11 | Sensor (Body/Tubes) GA Drawing | Design & Engineering | All | Face-to-face, tube material, pressure/temperature class, secondary containment rating/venting | R | Yes |  | Yes | Yes | Yes
12 | Transmitter/Converter GA & Wiring Diagram | Design & Engineering | All | Terminals, entries/glands, enclosure rating, output boards | R | Yes |  | Yes | Yes | Yes
13 | Remote Sensor–Transmitter Cable Specification (if remote) | Design & Engineering | If remote mount | Cable types/cores/shielding, max length, routing, part numbers | R | If applicable |  | If applicable | If applicable | Yes
14 | Materials of Construction List (wetted & non-wetted) | Design & Engineering | All | Tubes, manifolds, process connections, gaskets, housing | R | Yes |  | Yes | Yes | Yes
15 | Entrained Gas / Two-Phase Flow Handling | Design & Engineering | If gas in liquid likely | Diagnostics/drive-gain limits, EGM/EGP handling, alarms | R | If applicable |  | If applicable | If applicable | Yes
16 | Compliance Matrix to Project Specs / Datasheet | Design & Engineering | All | Point-by-point compliance; list deviations | R | Yes |  | Yes | Yes | Yes
17 | Custody Transfer Conformity (if CT service) | Design & Engineering | If CT required | API MPMS 5.6 (liquids), AGA Report No. 11 (gas), MID/MI-005, OIML R117 as applicable | R | If applicable |  | Yes | If applicable | Yes
18 | Factory Mass Flow Calibration Procedure | Procedures & QC | All | ISO/IEC 17025 traceability; points across range; meter factor if applicable | R | Yes |  | No | Yes | Yes
19 | Factory Density Calibration Procedure | Procedures & QC | All | Reference fluids, temp compensation method | R | Yes |  | No | Yes | Yes
20 | Zero Verification/Adjustment Procedure | Procedures & QC | All | Stabilization time, acceptance criteria | R | Yes |  | No | Yes | Yes
21 | Functional Test Procedure | Procedures & QC | All | Outputs (4–20 mA/pulse/frequency), totalizers, diagnostics, alarms | R | Yes |  | No | Yes | Yes
22 | Hydrostatic/Leak Test Procedure — sensor body | Procedures & QC | All sensors | Test pressure, duration, acceptance criteria | R/W | Yes |  | No | Yes | Yes
23 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Shipping caps, humidity/shock indicators if specified | R | Yes |  | If requested | Yes | Yes
24 | Software/Firmware Handling & Cybersecurity Hardening Guide | Procedures & QC | Smart transmitters | Default passwords, write lock, security switch, OTA policy | R | Yes |  | If requested | Yes | Yes
25 | Oxygen Service Cleaning Procedure | Procedures & QC | If O2 service | Cleaning verification method & acceptance | R | If applicable |  | No | If applicable | Yes
26 | Hazardous Area Certificates (ATEX/IECEx) + Schedules | Compliance & Certifications | If Ex required | Entity parameters, temp class, EPL/Zone, install conditions | R/F | If applicable |  | If applicable | If applicable | Yes
27 | EMC/EMI Compliance Statement | Compliance & Certifications | All | IEC 61326 series or equivalent | F | Yes |  | Yes | Yes | Yes
28 | Electrical Safety (IEC 61010-1) Declaration | Compliance & Certifications | All | Safety of electrical equipment | F | Yes |  | Yes | Yes | Yes
29 | Ingress Protection (IP66/67) Compliance Statement | Compliance & Certifications | All | IEC 60529 | F | Yes |  | Yes | Yes | Yes
30 | Conformity to ISO 10790 / ASME MFC-11M (Coriolis) | Compliance & Certifications | All | Design/accuracy/test basis per standard | F | Yes |  | Yes | Yes | Yes
31 | PED/UKCA/CE Declaration (as applicable) | Compliance & Certifications | Jurisdiction dependent | Category/Module or SEP | F | If applicable |  | If applicable | If applicable | Yes
32 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Tubes/manifolds/fasteners (wetted) | F | If applicable |  | If applicable | If applicable | Yes
33 | RoHS/REACH or regional material compliance | Compliance & Certifications | If contract requires |  | F | If applicable |  | If applicable | If applicable | Yes
34 | 3-A/EHEDG Hygienic Approval | Compliance & Certifications | If sanitary service | Hygienic design certificate/schedule | F | If applicable |  | If applicable | If applicable | Yes
35 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | All wetted metallics | Tubes, manifolds, process connections, fasteners | F | Yes |  | No | Yes | Yes
36 | Elastomer/Polymer Compliance (FDA/WRAS if required) | Materials & Special Service | If specified | O-rings/gaskets approvals | F | If applicable |  | If applicable | If applicable | Yes
37 | Mass Flow Calibration Certificate | Inspection & Test Records | All | Method (gravimetric/volumetric), points across range, expanded uncertainty | F/W | Yes |  | No | Yes | Yes
38 | Density Calibration Certificate | Inspection & Test Records | All | Reference fluids, temperature compensation | F/W | Yes |  | No | Yes | Yes
39 | Zero Verification Record (as-left) | Inspection & Test Records | All | Zero stability value and acceptance | F | Yes |  | No | Yes | Yes
40 | Functional Test Report | Inspection & Test Records | All | Outputs, diagnostics (drive gain/frequency), totalizers | F | Yes |  | No | Yes | Yes
41 | Hydrostatic/Pressure Test Certificate — sensor body | Inspection & Test Records | All sensors | Pressure, duration, acceptance | F/W | Yes |  | No | Yes | Yes
42 | Entrained Gas / Two-Phase Test Record | Inspection & Test Records | If specified | Alarm/diagnostic behavior under two-phase | F | If applicable |  | No | If applicable | Yes
43 | Environmental/Temperature Soak Test Record (if specified) | Inspection & Test Records | If contract requires | Hot/cold soak, output stability | F | If applicable |  | No | If applicable | Yes
44 | Vibration Test Record (if specified) | Inspection & Test Records | If contract requires | Frequency/acceleration profile | F | If applicable |  | No | If applicable | Yes
45 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes |  | No | Yes | Yes
46 | EDD/DTM/Device Files & Version List | Software & Integration | Smart devices | Provide files and version compatibility matrix | F | Yes |  | No | Yes | Yes
47 | GSD/EDS/CFH files (Profibus PA/FF/Modbus) | Software & Integration | If using FF/PA/Modbus | Node defaults, function blocks | F | If applicable |  | If applicable | If applicable | Yes
48 | Communication Conformance Certificate (HART/FF/PA) | Software & Integration | Digital comms devices | FieldComm Group/PI conformance | R/F | If applicable |  | If applicable | If applicable | Yes
49 | Host System Integration Notes | Software & Integration | All | DCS/AMS/PRM/800xA/FDM compatibility and parameters | F | Yes |  | If requested | Yes | Yes
50 | Firmware Revision Report & Release Notes | Software & Integration | All | Device firmware version per serial; issues/mitigations | F | Yes |  | No | Yes | Yes
51 | Modbus/HART Variable Map & Scaling | Software & Integration | If Modbus/HART used | Variable list, units, scaling, status bits | F | If applicable |  | If applicable | If applicable | Yes
52 | Installation, Operation & Maintenance (IOM) Manual | Manuals & Installation | All | Orientation/support, stress isolation, purging/cleaning, grounding | F | Yes |  | If requested | Yes | Yes
53 | Safety Manual (IEC 61508) | Manuals & Installation | If used in SIS | Proof test intervals, failure rates, restrictions | F | If applicable |  | If applicable | If applicable | Yes
54 | Mounting/Bracket Drawings & Torque Values | Manuals & Installation | All | Bolting/torque tables, gasket guidance | F | Yes |  | If requested | Yes | Yes
55 | Storage & Handling Instructions | Manuals & Installation | All | Caps, humidity/temp storage, protection of tubes/sensors | F | Yes |  | No | Yes | Yes
56 | As-built Configuration & Parameter List (per tag) | As-built & Close-out | All | Ranges (mass/density/temp), outputs, damping/filters, comm address, totalizer setup | F | Yes |  | No | Yes | Yes
57 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include firmware version and model code | F | Yes |  | No | Yes | Yes
58 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files as applicable | F | Yes |  | No | Yes | Yes
59 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes |  | No | Yes | Yes
60 | Preservation & Packing Records with Photos | Packing & Logistics | All | Transit shock/humidity indicators if specified | F | Yes |  | No | Yes | Yes
61 | Shipping, Handling & Lifting Instructions | Packing & Logistics |  | Orientation, lifting points, tube protection, secondary containment venting | F | Yes | All | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — Coriolis Flowmeter VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the service/scope requires (e.g., Ex area, custody transfer, oxygen service).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: Datasheet, Application/sizing check, Mass & density accuracy, Zero stability, Sensor & Converter GA, Materials list, Compliance matrix, ISO 10790/ASME MFC-11M conformity, EMC/Safety/IP declarations.
4) If custody transfer: include API MPMS 5.6 (liquids) / AGA 11 (gas) conformity plan and calibration approach; consider MID/MI-005 or OIML R117 as applicable.
5) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/config docs — 3–4 weeks; Final records — prior to shipment.
6) MRB content should include: mass & density calibration certificates, zero verification, functional test records, conformance certificates, materials certificates, manuals, as-built configuration, and packing list.