## Displacer Level Transmitter VDR
Vendor Document Requirement (VDR) — Displacer Level Transmitter — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration & witness points | Project Admin & Planning | All | Include chamber hydrotest, leak test, NDE hold points, calibration/functional tests | R/W | Yes | If requested | Yes | Yes
4 | Manufacturing / Assembly Schedule | Project Admin & Planning | All | Milestones and document submittal plan | R | Yes | No | Yes | No
5 | Nonconformance/Concession/Deviation & Corrective Action templates | Project Admin & Planning | All |  | R | Yes | No | Yes | No
6 | Instrument Datasheet — Displacer Level Transmitter (per tag) | Design & Engineering | All | Service (level/interface), range, SG range, temp/pressure, mounting, outputs/comms, Ex rating | R | Yes | Yes | Yes | Yes
7 | Application & Sizing Check | Design & Engineering | All | Fluid SGs (min/nom/max), temp/pressure window, min measurable Δbuoyancy, displacer length selection, turbulence/bypass needs | R | Yes | Yes | Yes | Yes
8 | Accuracy & Performance Statement | Design & Engineering | All | Reference accuracy, linearity, repeatability, temp/SG effect, hysteresis | R | Yes | Yes | Yes | Yes
9 | Displacer & Chamber GA Drawing | Design & Engineering | If chamber supplied | Face-to-face, nozzle orientation, vents/drains, guides, support brackets, C-to-C dimensions | R | If applicable | Yes | If applicable | Yes
10 | Top-mounted Displacer GA Drawing | Design & Engineering | If top-mounted | Nozzle details, insertion length, guide wire/rod, travel stops | R | If applicable | Yes | If applicable | Yes
11 | Transmitter/Converter GA & Wiring Diagram | Design & Engineering | All | Terminals, jumpers/switches, enclosure rating/entries | R | Yes | Yes | Yes | Yes
12 | Materials of Construction List (wetted & non-wetted) | Design & Engineering | All | Displacer, stem/spring, torque tube/lever, chamber/bridles, gaskets, fasteners, housing | R | Yes | Yes | Yes | Yes
13 | Pressure/Temperature Derating Curves — Chamber/Connections | Design & Engineering | If chamber supplied | Allowable service envelope, flange rating | R | If applicable | Yes | If applicable | Yes
14 | Buoyancy & Trim Calculation Sheet | Design & Engineering | All | Displacer volume/weight, SG at LRV/URV, spring trim/zero/span, interface calculation if applicable | R | Yes | Yes | Yes | Yes
15 | Interface Measurement Strategy (if interface) | Design & Engineering | If interface service | SG1/SG2 values, temp compensation, calibration method | R | If applicable | If applicable | If applicable | Yes
16 | Compliance Matrix to Project Specs / Datasheet | Design & Engineering | All | Point-by-point compliance; list deviations | R | Yes | Yes | Yes | Yes
17 | Calibration Procedure — Displacer (wet/weights) | Procedures & QC | All | Use of test fluid or calibrated weights; as-left LRV/URV; SG compensation | R | Yes | No | Yes | Yes
18 | Functional Test Procedure | Procedures & QC | All | Output scaling (4–20 mA/HART etc.), simulation, damping/filters, failure modes | R | Yes | No | Yes | Yes
19 | Hydrostatic/Leak Test Procedure — Chamber/Assembly | Procedures & QC | If chamber supplied | Test pressure, duration, acceptance criteria | R/W | If applicable | If applicable | If applicable | Yes
20 | WPS/PQR/WPQR — Chamber Fabrication | Procedures & QC | If fabricated chamber | Procedure and welder qualifications | R | If applicable | If applicable | If applicable | Yes
21 | NDE Procedures (VT/PT/RT/UT) — Chamber Welds | Procedures & QC | If fabricated chamber | Coverage and acceptance criteria | R | If applicable | If applicable | If applicable | Yes
22 | PMI Procedure for Alloy Verification | Procedures & QC | If specified | Materials to be verified and acceptance | R | If applicable | If applicable | If applicable | Yes
23 | Painting/Coating Procedure & DFT Plan | Procedures & QC | If painting/coating required | System, surface prep, SSPC/NACE class, DFT gauges & hold points | R | If applicable | If applicable | If applicable | Yes
24 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Shipping locks for displacer, moisture protection, caps | R | Yes | No | Yes | Yes
25 | Software/Firmware Handling & Cybersecurity Guide | Procedures & QC | Smart transmitters | Default passwords, write lock/security switch, update policy | R | If applicable | If applicable | If applicable | Yes
26 | Hazardous Area Certificates (ATEX/IECEx) + Schedules | Compliance & Certifications | If Ex required | Entity parameters, temp class, EPL/Zone, install conditions | R/F | If applicable | Yes | If applicable | Yes
27 | EMC/EMI Compliance Statement | Compliance & Certifications | If electronic transmitter | IEC 61326 series | F | If applicable | If applicable | If applicable | Yes
28 | Electrical Safety (IEC 61010-1) Declaration | Compliance & Certifications | If electronic transmitter | Safety of electrical equipment | F | If applicable | If applicable | If applicable | Yes
29 | Ingress Protection (IP66/67) Compliance Statement | Compliance & Certifications | If enclosure present | IEC 60529 | F | If applicable | If applicable | If applicable | Yes
30 | PED/UKCA/CE Declaration (as applicable) | Compliance & Certifications | Chamber as pressure accessory | Category/Module or SEP | F | If applicable | Yes | If applicable | Yes
31 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Displacer/chamber/trim (wetted) | F | If applicable | If applicable | If applicable | Yes
32 | Functional Safety (IEC 61508) — SIL Certificate & Safety Manual | Compliance & Certifications | If used in SIS | SIL capability, SFF, PFDavg assumptions | R/F | If applicable | If applicable | If applicable | Yes
33 | RoHS/REACH or regional material compliance | Compliance & Certifications | If contract requires |  | F | If applicable | If applicable | If applicable | Yes
34 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | All wetted metallics | Chamber body/nozzles, displacer, springs, fasteners | F | Yes | If applicable | Yes | Yes
35 | Elastomer/Polymer Compliance (FDA/WRAS if required) | Materials & Special Service | If specified | Gaskets/seals approvals | F | If applicable | If applicable | If applicable | Yes
36 | Calibration Certificate (as-left) — Level/Interface | Inspection & Test Records | All | Points across span; simulated SGs or test fluid; expanded uncertainty | F/W | Yes | No | Yes | Yes
37 | Functional Test Report | Inspection & Test Records | All | Output scaling, damping/filters, failure modes | F | Yes | No | Yes | Yes
38 | Hydrostatic/Pressure Test Certificate — Chamber/Assembly | Inspection & Test Records | If chamber supplied | Pressure, duration, acceptance | F/W | If applicable | No | If applicable | Yes
39 | Leak Test / Helium Tightness Test Record | Inspection & Test Records | If specified | Measured rate vs acceptance | F | If applicable | If applicable | If applicable | Yes
40 | NDE Reports (VT/PT/RT/UT) — Chamber Welds | Inspection & Test Records | If fabricated chamber | Coverage & results | F | If applicable | If applicable | If applicable | Yes
41 | Dimensional Inspection Report — Displacer & Chamber | Inspection & Test Records | All/if chamber | Displacer OD/length/weight; chamber ID/face-to-face; nozzle C-to-C; flange details | F | Yes | No | Yes | Yes
42 | Painting/Coating DFT Report | Inspection & Test Records | If painted/coated | DFT readings vs spec | F | If applicable | If applicable | If applicable | Yes
43 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes | No | Yes | Yes
44 | EDD/DTM/Device Files & Version List | Software & Integration | Smart devices | Provide files and compatibility matrix | F | If applicable | If applicable | If applicable | Yes
45 | Communication Conformance Certificate (HART/FF/PA) | Software & Integration | Digital comms devices | FieldComm Group/PI conformance | R/F | If applicable | If applicable | If applicable | Yes
46 | Host System Integration Notes | Software & Integration | All | DCS/AMS/PRM/800xA/FDM compatibility and parameters | F | Yes | If requested | Yes | Yes
47 | Firmware Revision Report & Release Notes | Software & Integration | All smart devices | Device firmware per serial; issues/mitigations | F | If applicable | If applicable | If applicable | Yes
48 | Installation, Operation & Maintenance (IOM) Manual | Manuals & Installation | All | Mounting, displacer shipping lock removal, zero/span adjustment, interface setup, wiring | F | Yes | If requested | Yes | Yes
49 | Storage & Handling Instructions | Manuals & Installation | All | Protection for displacer, moisture control, lifting/handling | F | Yes | No | Yes | Yes
50 | As-built Configuration & Parameter List (per tag) | As-built & Close-out | All | Range, units, damping, outputs, comm address, trim factors; SG values used | F | Yes | No | Yes | Yes
51 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include model code and firmware (if any) | F | Yes | No | Yes | Yes
52 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files as applicable | F | Yes | No | Yes | Yes
53 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes | No | Yes | Yes
54 | Preservation & Packing Records with Photos | Packing & Logistics | All | Shipping locks fitted; caps; shock/humidity indicators if specified | F | Yes | No | Yes | Yes
55 | Shipping & Handling Instructions | Packing & Logistics | All | Orientation, lifting points, protection of displacer/lever mechanisms | F | Yes | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — Displacer Level Transmitter VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the service/scope requires (e.g., chamber supplied, Ex area, sour service, SIS use, painting).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: Datasheet, Application & sizing check, Accuracy statement, GA drawings (displacer/chamber or top-mount), Materials list, Buoyancy/trim calc, Pressure/temperature rating, Compliance matrix, and (if applicable) Ex/IP/EMC/Electrical safety declarations.
4) If interface measurement: define SG1/SG2 and calibration method (wet/weights); specify operating temperature and impact on SG.
5) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/calculation docs — 3–4 weeks; Final records — prior to shipment.
6) MRB content should include: calibration & functional test records, hydro/leak test and NDE (if chamber), conformance certificates (Ex/IP/EMC/PED/NACE as applicable), materials certificates, manuals, as-built configuration, and packing list.