## Float & Tape Level VDR
Vendor Document Requirement (VDR) — Float & Tape Level Transmitter — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Applicibility | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes |  | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes |  | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration & witness points | Project Admin & Planning | All | Include travel/linearization verification, output checks, Ex/IP verification; hydro/NDE only if chamber/bridle supplied | R/W | Yes |  | If requested | Yes | Yes
4 | Manufacturing / Assembly Schedule | Project Admin & Planning | All | Milestones and document submittal plan | R | Yes |  | No | Yes | No
5 | Nonconformance/Concession/Deviation & Corrective Action templates | Project Admin & Planning | All |  | R | Yes |  | No | Yes | No
6 | Instrument Datasheet — Float & Tape Level Transmitter (per tag) | Design & Engineering | All | Measured variable (level/interface if used), range/travel, tank height, process temp/pressure, mounting nozzle, outputs/comms, Ex rating | R | Yes |  | Yes | Yes | Yes
7 | Application & Sizing Check | Design & Engineering | All | Fluid SG range, turbulence/slosh, foam/vapor effects, stilling well/guide requirement, min/max travel & blind zones | R | Yes |  | Yes | Yes | Yes
8 | Accuracy, Linearity & Hysteresis Statement | Design & Engineering | All | Reference accuracy, repeatability, mechanical hysteresis, temperature effects | R | Yes |  | Yes | Yes | Yes
9 | Transmitter Head/Converter GA & Wiring Diagram | Design & Engineering | All | Terminals, jumpers/switches, entry sizes, enclosure rating, local display/switch options | R | Yes |  | Yes | Yes | Yes
10 | Float Assembly GA Drawing | Design & Engineering | All | Float diameter/length/shape, material, guide rod/wire details, travel stops, attachment to tape/cable | R | Yes |  | Yes | Yes | Yes
11 | Tape/Cable & Pulley/Gearing Details | Design & Engineering | All | Tape/cable material & width/OD, corrosion resistance, pulley size, gear ratio, counterweight/tensioning method | R | Yes |  | Yes | Yes | Yes
12 | Stilling Well / Bridle / Guide Arrangement GA | Design & Engineering | If used | Nozzle details, supports, vent/drain, guide bars, internal baffles | R | If applicable |  | Yes | If applicable | Yes
13 | Materials of Construction List (wetted & non-wetted) | Design & Engineering | All | Float, tape/cable, guide, pulleys, housing, gaskets/elastomers, fasteners | R | Yes |  | Yes | Yes | Yes
14 | Pressure/Temperature Derating & Service Limits | Design & Engineering | All | Max process temp, pressure at nozzle/bridle, float material limits | R | Yes |  | Yes | Yes | Yes
15 | Conversion Ratio / Scaling Calculation Sheet | Design & Engineering | All | Float travel-to-output mapping, gearing ratio, linearization method, blind-zone allowance | R | Yes |  | Yes | Yes | Yes
16 | Compliance Matrix to Project Specs / Datasheet | Design & Engineering | All | Point-by-point compliance; list deviations | R | Yes |  | Yes | Yes | Yes
17 | Calibration & Linearization Procedure | Procedures & QC | All | Reference points across travel; correction/linearization method; expanded uncertainty | R | Yes |  | If requested | Yes | Yes
18 | Functional Test Procedure | Procedures & QC | All | Output 4–20 mA/HART/Modbus checks, alarms, NE43 upscale/downscale, diagnostics (NE107 if applicable) | R | Yes |  | If requested | Yes | Yes
19 | Mechanical Verification Procedure | Procedures & QC | All | Float buoyancy check, tape tension/counterweight set-up, pulley/gear integrity, friction/backlash check | R | Yes |  | No | Yes | Yes
20 | Hydrostatic/Leak Test Procedure — Bridle/Stilling well | Procedures & QC | If pressurized bridle/well | Test pressure, duration, acceptance criteria | R/W | If applicable |  | If applicable | If applicable | Yes
21 | WPS/PQR/WPQR — Bridle Fabrication | Procedures & QC | If fabricated bridle | Procedure and welder qualifications | R | If applicable |  | If applicable | If applicable | Yes
22 | NDE Procedures (VT/PT/RT/UT) — Bridle Welds | Procedures & QC | If fabricated bridle | Coverage and acceptance criteria | R | If applicable |  | If applicable | If applicable | Yes
23 | PMI Procedure for Alloy Verification | Procedures & QC | If specified | Wetted metallics to be verified | R | If applicable |  | If applicable | If applicable | Yes
24 | Painting/Coating Procedure & DFT Plan | Procedures & QC | If painting/coating required | System, surface prep, DFT gauges & hold points | R | If applicable |  | If applicable | If applicable | Yes
25 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Float & tape shipping locks, moisture barriers, caps; oil & grease-free if O₂ service | R | Yes |  | No | Yes | Yes
26 | Software/Firmware Handling & Cybersecurity Guide | Procedures & QC | If smart transmitter | Default passwords, write lock/security switch, firmware update policy | R | If applicable |  | No | If applicable | Yes
27 | Hazardous Area Certificates (ATEX/IECEx) + Schedules | Compliance & Certifications | If Ex required | Entity parameters, temp class, EPL/Zone, install conditions | R/F | If applicable |  | If applicable | If applicable | Yes
28 | EMC/EMI Compliance Statement (IEC 61326) | Compliance & Certifications | If electronic transmitter | Industrial environment | F | If applicable |  | If applicable | If applicable | Yes
29 | Electrical Safety (IEC 61010-1) Declaration | Compliance & Certifications | If electronic transmitter | Safety of electrical equipment | F | If applicable |  | If applicable | If applicable | Yes
30 | Ingress Protection (IP66/67) Compliance Statement | Compliance & Certifications | Transmitter head | IEC 60529 | F | Yes |  | Yes | Yes | Yes
31 | Functional Safety (IEC 61508) — SIL Certificate & Safety Manual | Compliance & Certifications | If used in SIS | SIL capability, SFF, PFDavg assumptions | R/F | If applicable |  | If applicable | If applicable | Yes
32 | PED/UKCA/CE Declaration (as applicable) | Compliance & Certifications | Bridle as pressure accessory | Category/Module or SEP | F | If applicable |  | If applicable | If applicable | Yes
33 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Float/bridle/fasteners (wetted) | F | If applicable |  | If applicable | If applicable | Yes
34 | RoHS/REACH or regional material compliance | Compliance & Certifications | If contract requires |  | F | If applicable |  | If applicable | If applicable | Yes
35 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | All wetted metallics | Float body, guides, bridle/well, fasteners | F | If applicable |  | If applicable | If applicable | Yes
36 | Elastomer/Polymer Compliance (FDA/WRAS if required) | Materials & Special Service | If specified | Gaskets/seals approvals | F | If applicable |  | If applicable | If applicable | Yes
37 | Calibration & Linearization Certificate (as-left) | Inspection & Test Records | All | Points across travel; error/linearity/hysteresis; expanded uncertainty | F/W | Yes |  | No | Yes | Yes
38 | Functional Test Report | Inspection & Test Records | All | Output scaling, alarms, diagnostics; NE43/NE107 where applicable | F | Yes |  | No | Yes | Yes
39 | Mechanical Verification Record | Inspection & Test Records | All | Float buoyancy, tape tension/counterweight set, pulley/gear inspection; travel length verification | F | Yes |  | No | Yes | Yes
40 | Hydrostatic/Pressure Test Certificate — Bridle/Stilling well | Inspection & Test Records | If pressurized bridle/well | Pressure, duration, acceptance | F/W | If applicable |  | No | If applicable | Yes
41 | NDE Reports (VT/PT/RT/UT) — Bridle Welds | Inspection & Test Records | If fabricated bridle | Coverage & results | F | If applicable |  | No | If applicable | Yes
42 | Painting/Coating DFT Report | Inspection & Test Records | If painted/coated | DFT readings vs spec | F | If applicable |  | No | If applicable | Yes
43 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes |  | No | Yes | Yes
44 | EDD/DTM/Device Files & Version List | Software & Integration | If smart transmitter | Provide files and compatibility matrix | F | If applicable |  | If applicable | If applicable | Yes
45 | Communication Conformance Certificate (HART/FF/PA) | Software & Integration | If digital comms | FieldComm Group/PI conformance | R/F | If applicable |  | If applicable | If applicable | Yes
46 | Modbus/HART Variable Map & Scaling | Software & Integration | If Modbus/HART used | Variables, units, scaling, status bits | F | If applicable |  | If applicable | If applicable | Yes
47 | Host System Integration Notes | Software & Integration | All | DCS/AMS/PRM/800xA/FDM compatibility and parameters | F | Yes |  | No | Yes | Yes
48 | Firmware Revision Report & Release Notes | Software & Integration | If smart transmitter | Firmware per serial; known issues/mitigations | F | If applicable |  | If applicable | If applicable | Yes
49 | Installation, Operation & Maintenance (IOM) Manual | Manuals & Installation | All | Mounting, float/tape routing, stilling/guide setup, linearization & scaling, wiring | F | Yes |  | If requested | Yes | Yes
50 | Storage & Handling Instructions | Manuals & Installation | All | Transport locks for float/tape, moisture control, corrosion protection | F | Yes |  | No | Yes | Yes
51 | As-built Configuration & Parameter List (per tag) | As-built & Close-out | All | Range, units, damping, linearization table, comm address, gear ratio used | F | Yes |  | No | Yes | Yes
52 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include model code and firmware (if any) | F | Yes |  | No | Yes | Yes
53 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files as applicable | F | Yes |  | No | Yes | Yes
54 | Packing List & Country of Origin (if required) | As-built & Close-out |  |  | F | Yes | All | No | Yes | Yes
55 | Preservation & Packing Records with Photos | Packing & Logistics | All | Float & tape locked/secured; shock/humidity indicators if specified | F | Yes |  | No | Yes | Yes
56 | Shipping & Handling Instructions | Packing & Logistics | All | Orientation, lifting/handling, re-commissioning steps after transport locks | F | Yes |  | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — Float & Tape Level Transmitter VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the service/scope requires (e.g., pressurized bridle/stilling well, Ex area, sour service, SIS use, smart transmitter).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: Datasheet, Application/sizing check, Accuracy/linearity/hysteresis statement, Head/Float/Tape GA drawings, Materials list, P/T service limits, Conversion/scaling calculations, Compliance matrix, and IP (and Ex/EMC where applicable).
4) If bridle or stilling well is in scope, include WPS/NDE/hydro plan and PED/CE as applicable. If smart transmitter head is included, provide EDD/DTM files and comms conformance.
5) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/config docs — 3–4 weeks; Final records — prior to shipment.
6) MRB content should include: calibration/linearization certificate, functional test, mechanical verification, hydro/NDE (if bridle), conformance certificates (Ex/IP/EMC/PED/NACE as applicable), materials certificates, manuals, as-built configuration, and packing list.