## GWR Level Transmitter VDR
Vendor Document Requirement (VDR) — Guided Wave Radar Level Transmitter — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration & witness points | Project Admin & Planning | All | Include functional test, echo-curve capture, chamber hydro/leak (if supplied), firmware lock checks | R/W | Yes | If requested | Yes | Yes
4 | Manufacturing / Assembly Schedule | Project Admin & Planning | All | Milestones and document submittal plan | R | Yes | No | Yes | No
5 | Nonconformance/Concession/Deviation & Corrective Action templates | Project Admin & Planning | All |  | R | Yes | No | Yes | No
6 | Instrument Datasheet — GWR Level Transmitter (per tag) | Design & Engineering | All | Service (level/interface), tank geometry, range, εr range, temp/pressure, mounting, outputs/comms, Ex rating | R | Yes | Yes | Yes | Yes
7 | Application & Sizing Check | Design & Engineering | All | Dielectric constant (εr) min/nom/max, foam/vapor/condensation effects, stilling well/bridle need, probe selection (rod/twin/coil/flexible) | R | Yes | Yes | Yes | Yes
8 | Accuracy & Performance Statement | Design & Engineering | All | Reference accuracy, linearity, repeatability, dead zone (blocking distance), near-zone performance | R | Yes | Yes | Yes | Yes
9 | Probe GA Drawing | Design & Engineering | All | Length (L), process connection, material, guides/weights, rigid vs flexible, insulation/coatings, trim rules | R | Yes | Yes | Yes | Yes
10 | Transmitter Head/Converter GA & Wiring Diagram | Design & Engineering | All | Terminals, jumpers/switches, entry sizes, enclosure rating, earthing/bonding | R | Yes | Yes | Yes | Yes
11 | Chamber/Bridle GA (if chamber/bridle supplied) | Design & Engineering | If chamber/bridle | Nozzles, vents/drains, guide bars, C-to-C dimensions, supports | R | If applicable | If applicable | If applicable | Yes
12 | Materials of Construction List (wetted & non-wetted) | Design & Engineering | All | Probe, process connection, chamber/bridle, gasket/elastomers, housing | R | Yes | Yes | Yes | Yes
13 | Pressure/Temperature Derating Curves — Probe/Chamber | Design & Engineering | All | Allowable service envelope, flange/thread rating, probe insulation temp limit | R | Yes | Yes | Yes | Yes
14 | Interface Measurement Strategy (if oil/water etc.) | Design & Engineering | If interface service | Use of upper/lower εr, threshold/delta settings, reference values & validation method | R | If applicable | If applicable | If applicable | Yes
15 | False Echo Mapping / Obstruction Handling Plan | Design & Engineering | All | Nozzle/ladder/internal obstructions, echo suppression settings, baseline capture | R | Yes | Yes | Yes | Yes
16 | Compliance Matrix to Project Specs / Datasheet | Design & Engineering | All | Point-by-point compliance; list deviations | R | Yes | Yes | Yes | Yes
17 | Functional Test & Configuration Procedure | Procedures & QC | All | Simulation, address/range set, damping/filters, NAMUR NE43 fail-safe, NE107 diagnostics, echo-curve capture | R | Yes | If requested | Yes | Yes
18 | Factory Verification / Reference Check Procedure | Procedures & QC | All | Reference reflectors or internal simulator; probe length verification; interface mode verification (if applicable) | R | Yes | If requested | Yes | Yes
19 | Hydrostatic/Leak Test Procedure — Chamber/Assembly | Procedures & QC | If chamber/bridle supplied | Test pressure, duration, acceptance criteria | R/W | If applicable | If applicable | If applicable | Yes
20 | WPS/PQR/WPQR — Chamber Fabrication | Procedures & QC | If fabricated chamber | Procedure and welder qualifications | R | If applicable | If applicable | If applicable | Yes
21 | NDE Procedures (VT/PT/RT/UT) — Chamber Welds | Procedures & QC | If fabricated chamber | Coverage and acceptance criteria | R | If applicable | If applicable | If applicable | Yes
22 | PMI Procedure for Alloy Verification | Procedures & QC | If specified | Materials to be verified and acceptance | R | If applicable | If applicable | If applicable | Yes
23 | Painting/Coating Procedure & DFT Plan | Procedures & QC | If painting/coating required | System, surface prep, DFT gauges & hold points | R | If applicable | If applicable | If applicable | Yes
24 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Probe shipping protectors, coil winding protection, moisture barriers, caps | R | Yes | No | Yes | Yes
25 | Software/Firmware Handling & Cybersecurity Hardening Guide | Procedures & QC | Smart transmitters | Default passwords, write lock/security switch, update policy | R | Yes | No | Yes | Yes
26 | Hazardous Area Certificates (ATEX/IECEx) + Schedules | Compliance & Certifications | If Ex required | Entity parameters, temp class, EPL/Zone, install conditions | R/F | If applicable | If applicable | If applicable | Yes
27 | EMC/EMI Compliance Statement (IEC 61326) | Compliance & Certifications | All | Industrial environment | F | Yes | Yes | Yes | Yes
28 | Electrical Safety (IEC 61010-1) Declaration | Compliance & Certifications | All | Safety of electrical equipment | F | Yes | Yes | Yes | Yes
29 | Ingress Protection (IP66/67) Compliance Statement | Compliance & Certifications | All | IEC 60529 | F | Yes | Yes | Yes | Yes
30 | Functional Safety (IEC 61508) — SIL Certificate & Safety Manual | Compliance & Certifications | If used in SIS | SIL capability, SFF, PFDavg assumptions | R/F | If applicable | If applicable | If applicable | Yes
31 | PED/UKCA/CE Declaration (as applicable) | Compliance & Certifications | Chamber as pressure accessory | Category/Module or SEP | F | If applicable | If applicable | If applicable | Yes
32 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Probe/chamber/trim (wetted) | F | If applicable | If applicable | If applicable | Yes
33 | RoHS/REACH or regional material compliance | Compliance & Certifications | If contract requires |  | F | If applicable | If applicable | If applicable | Yes
34 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | All wetted metallics | Probe wetted parts, chamber body/nozzles, fasteners | F | Yes | If applicable | Yes | Yes
35 | Elastomer/Polymer Compliance (FDA/WRAS if required) | Materials & Special Service | If specified | Gaskets/seals approvals | F | If applicable | If applicable | If applicable | Yes
36 | Functional Test Report & Configuration As-left | Inspection & Test Records | All | Range, damping, outputs, address; NAMUR NE43/NE107 checks | F | Yes | No | Yes | Yes
37 | Echo Curve / Baseline Diagnostic Capture | Inspection & Test Records | All | Saved trace at shipment; false echo map if applied | F/W | Yes | No | Yes | Yes
38 | Probe Length & Dead-zone Verification Record | Inspection & Test Records | All | Measured L, blocking distance/near-zone, trim confirmation | F | Yes | No | Yes | Yes
39 | Interface Mode Verification Record | Inspection & Test Records | If interface service | εr thresholds, delta method, test results | F | If applicable | If applicable | If applicable | Yes
40 | Hydrostatic/Pressure Test Certificate — Chamber/Assembly | Inspection & Test Records | If chamber/bridle supplied | Pressure, duration, acceptance | F/W | If applicable | If applicable | If applicable | Yes
41 | NDE Reports (VT/PT/RT/UT) — Chamber Welds | Inspection & Test Records | If fabricated chamber | Coverage & results | F | If applicable | If applicable | If applicable | Yes
42 | Painting/Coating DFT Report | Inspection & Test Records | If painted/coated | DFT readings vs spec | F | If applicable | If applicable | If applicable | Yes
43 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes | No | Yes | Yes
44 | EDD/DTM/Device Files & Version List | Software & Integration | Smart devices | Provide files and compatibility matrix | F | Yes | No | Yes | Yes
45 | Communication Conformance Certificate (HART/FF/PA) | Software & Integration | Digital comms devices | FieldComm Group/PI conformance | R/F | If applicable | If applicable | If applicable | Yes
46 | Modbus/HART Variable Map & Scaling | Software & Integration | If Modbus/HART used | Variables, units, status bits, echo-curve export support | F | If applicable | If applicable | If applicable | Yes
47 | Host System Integration Notes | Software & Integration | All | DCS/AMS/PRM/800xA/FDM compatibility and parameters | F | Yes | No | Yes | Yes
48 | Firmware Revision Report & Release Notes | Software & Integration | All smart devices | Device firmware per serial; known issues/mitigations | F | Yes | No | Yes | Yes
49 | Installation, Operation & Maintenance (IOM) Manual | Manuals & Installation | All | Mounting, blocking distance/near-zone, nozzle effects, false echo mapping, interface setup, wiring | F | Yes | If requested | Yes | Yes
50 | Storage & Handling Instructions | Manuals & Installation | All | Probe transport protection, coil winding handling (flexible), moisture control | F | Yes | No | Yes | Yes
51 | As-built Configuration & Parameter List (per tag) | As-built & Close-out | All | Range, units, damping, blocking distance, εr thresholds, interface settings, comm address | F | Yes | No | Yes | Yes
52 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include model code and firmware version | F | Yes | No | Yes | Yes
53 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files as applicable | F | Yes | No | Yes | Yes
54 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes | No | Yes | Yes
55 | Preservation & Packing Records with Photos | Packing & Logistics | All | Probe protectors, coil secured, caps, shock/humidity indicators if specified | F | Yes | No | Yes | Yes
56 | Shipping & Handling Instructions | Packing & Logistics | All | Orientation, lifting/handling, flexible probe unwinding guidance | F | Yes | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — GWR Level Transmitter VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the service/scope requires (e.g., chamber/bridle supplied, Ex area, sour service, SIS use).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: Datasheet, Application & sizing check (εr/foam/vapor), Accuracy statement, Probe & Head GA, Materials list, P/T derating, False-echo mapping plan, Compliance matrix, and EMC/Safety/IP declarations; Ex/SIL as applicable.
4) Interface measurement: provide εr thresholds/delta method and verification plan. For chamber installs, include chamber GA and hydro/NDE plan.
5) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/config docs — 3–4 weeks; Final records — prior to shipment.
6) MRB content should include: functional test & echo curve, probe length/dead-zone verification, interface verification (if any), hydro/NDE for chamber, conformance certificates (Ex/IP/EMC/PED/NACE as applicable), materials certificates, manuals, as-built configuration, and packing list.