## GC VDR
Vendor Document Requirement (VDR) — Gas Chromatograph (GC) — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Aligned with ISO 9001 supplier QMS | R | Yes | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. FAT/iFAT, witness & hold points | Project Admin & Planning | All | Performance tests, validation, leak/pressure tests, software tests | R/W | Yes | If requested | Yes | Yes
4 | Manufacturing / Assembly / Software Build Schedule | Project Admin & Planning | All | Milestones and submittal plan | R | Yes | No | Yes | No
5 | Deviation/Concession (NCR) & Corrective Action templates | Project Admin & Planning | All |  | R | Yes | No | Yes | No
6 | Analyzer Datasheet — Process GC (per tag) | Design & Engineering | All | Stream count, oven type, detectors (TCD/FID/FPD/PDHID), columns, cycle time, Ex rating, outputs/comms | R | Yes | Yes | Yes | Yes
7 | Application & Composition Basis | Design & Engineering | All | Feed/stream compositions, expected ranges/limits, dew point & condensation risk, contaminants (H2S, moisture, solids) | R | Yes | Yes | Yes | Yes
8 | Method Description & Component List | Design & Engineering | All | Target components, ranges, detection limits, internal standards, cut times/valve timings | R | Yes | Yes | Yes | Yes
9 | Column Set & Detector Selection Rationale | Design & Engineering | All | Column type/length/ID/phase; detector choice; separation & sensitivity | R | Yes | Yes | Yes | Yes
10 | Cycle Time & Availability Statement | Design & Engineering | All | Analysis time, stream switching intervals, duty/standby philosophy | R | Yes | Yes | Yes | Yes
11 | Sample System P&ID / Hook-up (per analyzer) | Design & Engineering | All | Probe, filters, regulators, fast loops, heaters/heat tracing, bypass/vent, validation/calibration gas routing | R | Yes | Yes | Yes | Yes
12 | Sample Probe & Extraction Details | Design & Engineering | All | Insertion length, location, velocity profile, API 14.1/ISO 10715 guidance where applicable | R | Yes | Yes | Yes | Yes
13 | Heated Enclosures/Lines & Heat Load Calculation | Design & Engineering | If heating used | Heated line specs, setpoints, load/IB currents, insulation; moisture/hydrocarbon dew point control | R | If applicable | If applicable | If applicable | Yes
14 | Carrier/Utility Gas Specifications | Design & Engineering | All | Carrier, fuel, zero/span/validation gases; purity; dew point; pressure & flow; cylinder manifolds; ISO 8573-1 air quality (if pneumatic) | R | Yes | Yes | Yes | Yes
15 | Vent/Exhaust & Safe Dispersion Plan | Design & Engineering | All | Vent location, backpressure, flame arrestors (if required), environmental constraints | R | Yes | Yes | Yes | Yes
16 | Bill of Materials — Analyzer + Sample System | Design & Engineering | All | Part nos., materials, pressure ratings, wetted/non-wetted list | R | Yes | Yes | Yes | Yes
17 | Compliance Matrix to Project Specs | Design & Engineering | All | Point-by-point compliance; deviations listed | R | Yes | Yes | Yes | Yes
18 | Accuracy, Repeatability, Linearity & Drift Statement | Analytical Performance | All | Per component; ambient & utility variations considered | R | Yes | Yes | Yes | Yes
19 | Retention Times, Windows & Integration Parameters | Analytical Performance | All | Chromatogram gating, baseline methods, alarms for RT drift | R | Yes | Yes | Yes | Yes
20 | Natural Gas Property Calculations (if applicable) | Analytical Performance | Nat. gas/energy service | Select: ISO 6976 (HHV/LHV), AGA-8 (compressibility), GPA 2172 (HV at base); speed of sound (AGA-10) if required | R | If applicable | Yes | If applicable | Yes
21 | Validation/Calibration Strategy | Analytical Performance | All | Auto-cal/auto-val frequency, criteria, out-of-tolerance actions, audit trail | R | Yes | Yes | Yes | Yes
22 | Analyzer Shelter/House GA & Layout | Utilities/Shelter/HAZ | If shelter supplied | Rack layouts, HVAC, gas detection, lighting, doors, purge panels | R | If applicable | Yes | If applicable | Yes
23 | Shelter HVAC & Pressurization/Purge Calculations | Utilities/Shelter/HAZ | If shelter/Ex p used | IEC 60079-2 or NFPA 496; leak rates; purge time; differential pressure & airflow | R | If applicable | Yes | If applicable | Yes
24 | Hazardous Area Classification & Device Selection | Utilities/Shelter/HAZ | If Ex area | ATEX/IECEx marking; EPL/Zone; T-class; install conditions | R | If applicable | Yes | If applicable | Yes
25 | Power Distribution & Load List | Utilities/Shelter/HAZ | All | UPS feeders, PSUs, MCBs, heater loads, redundancy | R | Yes | Yes | Yes | Yes
26 | Earthing, Bonding & Shielding Scheme | Utilities/Shelter/HAZ | All | PE/FE bars, IS barriers/isolators, shield terminations | R | Yes | Yes | Yes | Yes
27 | Gas Detection & ESD/Shutdown Interlocks | Utilities/Shelter/HAZ | If shelter/Ex area | Detector placement, setpoints, interlocks to purge/HVAC/ESD | R | If applicable | Yes | If applicable | Yes
28 | Factory Performance Test Procedure (FAT) — GC | Procedures & QC | All | Accuracy/linearity/repeatability; validation run; chromatograms capture; alarms/diagnostics | R/W | Yes | If requested | Yes | Yes
29 | Sample System Leak/Pressure/Heater Test Procedure | Procedures & QC | All sample systems | Leak rates, pressure tests, heater setpoints & trips | R/W | Yes | If requested | Yes | Yes
30 | Calibration/Validation Gas Handling & Safety Procedure | Procedures & QC | All | Cylinders, regulators, purge/vent, expiry tracking | R | Yes | If requested | Yes | Yes
31 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Protect columns/detectors; caps; desiccants; transport locks | R | Yes | No | Yes | Yes
32 | ATEX/IECEx Certificates + Schedules (Analyzer, Purge panel, IS barriers) | Compliance & Certifications | If Ex required | Entity parameters, temp class, EPL/Zone, install conditions | R/F | If applicable | Yes | If applicable | Yes
33 | EMC/EMI Compliance Statement (IEC 61326) | Compliance & Certifications | All | Industrial environment | F | Yes | Yes | Yes | Yes
34 | Electrical Safety (IEC 61010-1) Declaration | Compliance & Certifications | All | Safety of electrical equipment | F | Yes | Yes | Yes | Yes
35 | Ingress Protection (IP/NEMA) Compliance Statement | Compliance & Certifications | Enclosures/shelters | IEC 60529; NEMA where applicable | F | Yes | Yes | Yes | Yes
36 | PED/CE/UKCA Declaration (as applicable) | Compliance & Certifications | Pressure accessories (sample system) | Category/Module or SEP | F | If applicable | If applicable | If applicable | Yes
37 | Natural Gas Analytical Standards Declaration | Compliance & Certifications | Nat. gas service | ISO 6974/6976, AGA-8/10, GPA 2261/2172 as applicable | F | If applicable | If applicable | If applicable | Yes
38 | CEMS/Environmental Compliance (if scope) | Compliance & Certifications | If CEMS | EN 15267 (QAL1 listing), EN 14181 QAL2/QAL3 procedure | R/F | If applicable | If applicable | If applicable | Yes
39 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | Sample system wetted | Probes, regulators, filters, tubing, manifolds | F | If applicable | If applicable | If applicable | Yes
40 | Elastomer/Polymer Compliance (FDA/WRAS if required) | Materials & Special Service | If hygienic or specified | Seals/O-rings approvals | F | If applicable | If applicable | If applicable | Yes
41 | FAT Performance Test Report — GC | Inspection & Test Records | All | Accuracy/linearity/repeatability; chromatograms; alarms/diagnostics logs | F/W | Yes | No | Yes | Yes
42 | Validation Run Report | Inspection & Test Records | All | Validation gas results vs acceptance; drift since last cal | F | Yes | No | Yes | Yes
43 | Chromatogram Set (PDF/native) — Method Baseline | Inspection & Test Records | All | RT windows & integration marks for each component | F | Yes | No | Yes | Yes
44 | Sample System Leak/Pressure/Heater Test Certificates | Inspection & Test Records | All | Leak rates, pressure tests, heater setpoints | F/W | Yes | No | Yes | Yes
45 | Calibration/Validation Gas Certificates (ISO/IEC 17025 traceable) | Inspection & Test Records | All | Blend tolerances, expiry, lot/batch numbers | F | Yes | No | Yes | Yes
46 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes | No | Yes | Yes
47 | Analyzer System Architecture & Network Diagram | Software & Integration | All | Analyzer workstation/PC, servers, clients, remote access rules | R | Yes | If requested | Yes | Yes
48 | Data Map & Scaling — Modbus/OPC/Other | Software & Integration | All | Variables (components %, CV, RD, Wobbe, alarms), units & scaling; quality/status bits | F | Yes | No | Yes | Yes
49 | EDD/DTM/Viewer Software & License List | Software & Integration | All | Chromatogram viewer/editor; compliance with OS; license dongles/keys | F | Yes | No | Yes | Yes
50 | Firmware/Software Revision Report & Release Notes | Software & Integration | All | Controller/PC SW versions; known issues/mitigations | F | Yes | No | Yes | Yes
51 | Backup/Restore & Disaster Recovery Procedure | Software & Integration | All | Images, data directories, periodic export; verification tests | R | Yes | If requested | Yes | Yes
52 | Cybersecurity Hardening Checklist (Analyzer PCs) | Software & Integration | If Windows/Linux PC | Accounts, services, patching, whitelisting/AV, USB control, remote access | R | If applicable | If applicable | If applicable | Yes
53 | Installation, Operation & Maintenance (IOM) Manual — Analyzer & Sample System | Manuals & Installation | All | Install, start-up, method tuning, maintenance, troubleshooting | F | Yes | No | Yes | Yes
54 | Wiring Diagrams & Terminal Schedules | Manuals & Installation | All | Analyzer, shelter panels, junction boxes, alarms, heaters | F | Yes | No | Yes | Yes
55 | Mechanical GA & Mounting Templates | Manuals & Installation | All | Footprints, clearances, cable/pipe entries | F | Yes | No | Yes | Yes
56 | Storage & Handling Instructions | Manuals & Installation | All | Column/detector protection, temperature/moisture limits | F | Yes | No | Yes | Yes
57 | As-built Method, RT Windows & Integration Parameters | As-built & Close-out | All | Final method files; RT tables; integration config | F | Yes | No | Yes | Yes
58 | As-built Configuration & Parameter List (per analyzer) | As-built & Close-out | All | Streams, columns, detector settings, utilities setpoints, comms address | F | Yes | No | Yes | Yes
59 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Analyzer, PCs, purge panels, barriers; firmware/SW versions | F | Yes | No | Yes | Yes
60 | Recommended Spare Parts List & MTBF Data | As-built & Close-out | All | Columns, detectors, valves, filters, regulators, heated lines | F | Yes | No | Yes | Yes
61 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files as applicable | F | Yes | No | Yes | Yes
62 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes | No | Yes | Yes
63 | Preservation & Packing Records with Photos | Packing & Logistics | All | Transport locks, desiccants; shock/tilt indicators if specified | F | Yes | No | Yes | Yes
64 | Shipping & Handling Instructions | Packing & Logistics | All | Orientation, handling, storage, shipping marks | F | Yes | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — Gas Chromatograph (GC) VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the scope requires (e.g., Ex shelter with purge, natural gas energy calculations, CEMS compliance).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: analyzer datasheet; method & component list; column/detector rationale; cycle time & availability; sample system P&ID & probe details; utilities (carrier/validation gases) spec; vent plan; compliance matrix; EMC/Electrical/IP declarations; Ex certificates (if applicable).
4) Standards reference (typical): ISO 6974/6976, GPA 2172 & 2261, AGA 8/10 (natural gas properties); IEC 60079-2 (Ex p purge); IEC 61326 (EMC); IEC 61010-1 (safety); IEC 60529 (IP); NFPA 496 (purge) as applicable; EN 15267 & EN 14181 (CEMS).
5) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/calcs/procedures — 3–5 weeks; FAT/iFAT procedures — 4 weeks; Final records — prior to shipment.
6) MRB content should include: FAT performance & validation reports with chromatograms, calibration/validation gas certificates, conformance certificates (Ex/IP/EMC/PED as applicable), materials certificates for sample system, manuals, as-built method/configuration, and packing list.