## Pressure Gauge VDR
Vendor Document Requirement (VDR) — Pressure Gauge — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration & witness points | Project Admin & Planning | All | Include calibration, leak checks for seal assemblies; functional checks for contact gauges | R/W | Yes | If requested | Yes | Yes
4 | Manufacturing / Assembly Schedule | Project Admin & Planning | All | Milestones and submission plan | R | Yes | No | Yes | No
5 | Nonconformance/Concession/Deviation & Corrective Action templates | Project Admin & Planning | All |  | R | Yes | No | Yes | No
6 | Instrument Datasheet — Pressure Gauge (per tag) | Design & Engineering | All | Dial size, case type (solid front/blow-out back), connection, range, fill (dry/glycerin/silicone), window (glass/safety) | R | Yes | Yes | Yes | Yes
7 | Accuracy Class & Range/Overpressure Statement | Design & Engineering | All | EN 837-1 or ASME B40.100 accuracy class; overpressure/rupture disc details | R | Yes | Yes | Yes | Yes
8 | General Arrangement (GA) — Gauge | Design & Engineering | All | Mounting (bottom/back), panel/surface brackets, blow-out, restrictor/snubber | R | Yes | Yes | Yes | Yes
9 | Materials of Construction List (wetted & non-wetted) | Design & Engineering | All | Bourdon tube, socket, movement, case, window, gasket, fill plugs | R | Yes | Yes | Yes | Yes
10 | Diaphragm Seal/Capillary Assembly Datasheet | Design & Engineering | If seal/capillary supplied | Seal type (wafer/flanged/threaded), fill fluid (SDS), capillary length/armor, temperature limits | R | If applicable | If applicable | If applicable | Yes
11 | Process Connection Details & Restrictor/Orifice Spec | Design & Engineering | If specified | Thread/flange rating, restrictor bore for pulsation control | R | If applicable | If applicable | If applicable | Yes
12 | Ingress Protection (IP) & Case Venting/Draining | Design & Engineering | All | IEC 60529 IP rating; vent plug details for filled gauges | R | Yes | Yes | Yes | Yes
13 | Compliance Matrix to Project Specifications | Design & Engineering | All | Point-by-point compliance; deviations list | R | Yes | Yes | Yes | Yes
14 | Calibration/Verification Procedure | Procedures & QC | All | ISO/IEC 17025 traceability; up/down scale points; acceptable hysteresis | R | Yes | If requested | Yes | Yes
15 | Leak/Pressure Test Procedure — Seal Assemblies | Procedures & QC | If seal/cap supplied | Pressure, duration, acceptance | R/W | If applicable | If applicable | If applicable | Yes
16 | Helium/Vacuum Tightness Procedure | Procedures & QC | If vacuum service | Rate/limit and method | R | If applicable | If applicable | If applicable | Yes
17 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Transport plugs, dial/needle locks, shock indicators if specified | R | Yes | No | Yes | Yes
18 | Oxygen Service Cleaning Procedure | Procedures & QC | If O2 service | Cleaning method & verification | R | If applicable | If applicable | If applicable | Yes
19 | EN 837-1 / ASME B40.100 Conformity Statement | Compliance & Certifications | All | Accuracy class, safety pattern (S3/solid front) as applicable | F | Yes | Yes | Yes | Yes
20 | Ingress Protection (IEC 60529) Statement | Compliance & Certifications | All | IP rating achieved | F | Yes | Yes | Yes | Yes
21 | PED/CE/UKCA Declaration (as applicable) | Compliance & Certifications | Jurisdiction dependent | Pressure accessory category/module or SEP | F | If applicable | If applicable | If applicable | Yes
22 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Socket/diaphragm/wetted parts | F | If applicable | If applicable | If applicable | Yes
23 | RoHS/REACH or regional material compliance | Compliance & Certifications | If contract requires |  | F | If applicable | If applicable | If applicable | Yes
24 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | All wetted metallics | Socket, diaphragm, seals, capillary armor if wetted | F | If applicable | If applicable | If applicable | Yes
25 | Fill Fluid Safety Data Sheet (SDS) | Materials & Special Service | If filled/with seal | Glycerin/silicone/fluorinated fluids | F | If applicable | If applicable | If applicable | Yes
26 | Calibration Certificate (as-left) — Pressure Gauge | Inspection & Test Records | All | Points across scale; expanded uncertainty; reference instrument | F/W | Yes | No | Yes | Yes
27 | Leak/Pressure Test Certificate — Seal/Capillary | Inspection & Test Records | If seal/cap supplied | Pressure, duration, acceptance | F/W | If applicable | No | If applicable | Yes
28 | Helium/Vacuum Tightness Test Record | Inspection & Test Records | If vacuum service | Measured rate vs acceptance | F | If applicable | No | If applicable | Yes
29 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes | No | Yes | Yes
30 | Installation, Operation & Maintenance (IOM) Manual — Gauge/Seal | Manuals & Installation | All | Mounting orientation; venting of filled gauges; handling precautions | F | Yes | If requested | Yes | Yes
31 | Storage & Handling Instructions | Manuals & Installation | All | Shock/impact, temp limits for fill fluid | F | Yes | If requested | Yes | Yes
32 | As-built Tag List & Configuration (per tag) | As-built & Close-out | All | Dial size, range, connection, fill, seal/capillary details | F | Yes | No | Yes | Yes
33 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include model code | F | Yes | No | Yes | Yes
34 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files | F | Yes | No | Yes | Yes
35 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes | No | Yes | Yes
36 | Preservation & Packing Records with Photos | Packing & Logistics | All | Shock/humidity indicators if specified; dial/needle protection | F | Yes | No | Yes | Yes
37 | Shipping & Handling Instructions | Packing & Logistics | All | Orientation, lifting/handling cautions | F | Yes | No | Yes | Yes

## Temperature Gauge VDR
Vendor Document Requirement (VDR) — Temperature Gauge — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration & witness points | Project Admin & Planning | All | Include calibration; leak/pressure tests for thermowell; functional checks for adjustable angle | R/W | Yes | If requested | Yes | Yes
4 | Instrument Datasheet — Temperature Gauge (per tag) | Design & Engineering | All | Type (bimetal/gas), dial size, angle (axial/radial/adjustable), stem OD/length, range, connection, case, window | R | Yes | Yes | Yes | No
5 | Accuracy Class Statement | Design & Engineering | All | EN 13190 accuracy class; ambient influence | R | Yes | Yes | Yes | No
6 | General Arrangement (GA) — Gauge | Design & Engineering | All | Orientation, clamp/panel hardware, adjustable head details | R | Yes | Yes | Yes | Yes
7 | Thermowell Datasheet (if supplied) | Design & Engineering | If thermowell supplied | Type, material, insertion length U, tip shape; surface finish | R | If applicable | If applicable | If applicable | Yes
8 | Thermowell Wake Frequency & Stress Calculations | Design & Engineering | If thermowell supplied | ASME PTC 19.3 TW-2016 (or project spec) | R | If applicable | If applicable | If applicable | Yes
9 | Materials of Construction List (wetted & non-wetted) | Design & Engineering | All | Stem, head, dial, window, thermowell, seals | R | Yes | Yes | Yes | Yes
10 | Remote Capillary/Extension (if gas-actuated) | Design & Engineering | If remote | Capillary length, armor, ambient temp limits | R | If applicable | If applicable | If applicable | Yes
11 | Compliance Matrix to Project Specifications | Design & Engineering | All | Point-by-point compliance; deviations list | R | Yes | Yes | Yes | Yes
12 | Calibration/Verification Procedure | Procedures & QC | All | ISO/IEC 17025 traceability; points across range | R | Yes | If requested | Yes | Yes
13 | Hydro/Pressure Test Procedure — Thermowell | Procedures & QC | If thermowell supplied | Pressure, duration, acceptance | R/W | If applicable | If applicable | If applicable | Yes
14 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Stem/end caps, desiccant, dial protection | R | Yes | No | Yes | Yes
15 | EN 13190 Conformity Statement (Bourdon-type, bimetal, gas-actuated) | Compliance & Certifications | All | Accuracy class and type | F | Yes | Yes | Yes | Yes
16 | Ingress Protection (IEC 60529) Statement | Compliance & Certifications | All | IP rating achieved | F | Yes | Yes | Yes | Yes
17 | PED/CE/UKCA Declaration (as applicable) | Compliance & Certifications | If thermowell pressure accessory | Category/Module or SEP | F | If applicable | If applicable | If applicable | Yes
18 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Thermowell/stem materials | F | If applicable | If applicable | If applicable | Yes
19 | RoHS/REACH or regional material compliance | Compliance & Certifications | If contract requires |  | F | If applicable | If applicable | If applicable | Yes
20 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | Thermowell/stem | Bar/forging/tube certificates; fasteners | F | If applicable | If applicable | If applicable | Yes
21 | Calibration Certificate (as-left) — Temperature Gauge | Inspection & Test Records | All | Points across range; expanded uncertainty; reference instrument | F/W | Yes | No | Yes | Yes
22 | Dimensional Inspection Report — Stem/Thermowell | Inspection & Test Records | All/if supplied | OD/length, thread/flange details, clearances | F | Yes | No | Yes | Yes
23 | Hydro/Pressure Test Certificate — Thermowell | Inspection & Test Records | If required | Pressure, duration, acceptance | F/W | If applicable | No | If applicable | Yes
24 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes | No | Yes | Yes
25 | Installation, Operation & Maintenance (IOM) Manual — Gauge/Thermowell | Manuals & Installation | All | Insertion depth, angle adjustment, sealing/torque, thermal lag notes | F | Yes | If requested | Yes | Yes
26 | Storage & Handling Instructions | Manuals & Installation | All | Protection for stems/dials; transport packaging | F | Yes | If requested | Yes | Yes
27 | As-built Tag List & Configuration (per tag) | As-built & Close-out | All | Dial size, range, angle, stem OD/length, thermowell details | F | Yes | No | Yes | Yes
28 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include model code | F | Yes | No | Yes | Yes
29 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files | F | Yes | No | Yes | Yes
30 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes | No | Yes | Yes
31 | Preservation & Packing Records with Photos | Packing & Logistics | All | Dial/stem protection; humidity indicators if specified | F | Yes | No | Yes | Yes
32 | Shipping & Handling Instructions | Packing & Logistics | All | Orientation, handling cautions | F | Yes | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — Pressure & Temperature Gauges VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the service/scope requires (e.g., diaphragm seals/capillary, thermowells, sour service, vacuum or oxygen service).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: Datasheets, Accuracy/Performance statements, GA drawings, Materials lists, IP/PED/standard conformity, and Compliance matrices.
4) Standards reference (typical): EN 837-1 / ASME B40.100 (pressure gauges), EN 13190 (temperature gauges), IEC 60529 (IP), EN 10204 (MTC), NACE MR0175/ISO 15156 (sour). If thermowell: ASME PTC 19.3 TW-2016. If chemical seal: provide fill-fluid SDS.
5) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/config docs — 3–4 weeks; Final records — prior to shipment.
6) MRB content should include: calibration and test records, conformance certificates, materials certificates (incl. seal/thermowell), manuals, as-built lists, and packing list.