## Thermal Mass Flow VDR
Vendor Document Requirement (VDR) — Thermal Mass (Dispersion) Flowmeter — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration & witness points | Project Admin & Planning | All | Factory flow calibration, electronics checks, leak/pressure tests (inline), insertion hot-tap checks (if applicable) | R/W | Yes | If requested | Yes | Yes
4 | Manufacturing / Assembly Schedule | Project Admin & Planning | All | Milestones and document submittal plan | R | Yes | No | Yes | No
5 | Nonconformance/Concession/Deviation & Corrective Action templates | Project Admin & Planning | All |  | R | Yes | No | Yes | No
6 | Instrument Datasheet — Thermal Mass Flowmeter (per tag) | Design & Engineering | All | Service gas/mix, range, base conditions (Nm³/h or Sm³/h), pressure/temperature range, insertion vs in-line, outputs/comms, Ex rating | R | Yes | Yes | Yes | Yes
7 | Application & Sizing Check | Design & Engineering | All | Gas composition & variability, viscosity/density, pressure/temperature envelope, turndown, minimum velocity, straight-run/flow conditioner needs, moisture/particulates, installation orientation | R | Yes | Yes | Yes | Yes
8 | Accuracy, Repeatability & Response Statement | Design & Engineering | All | Stated accuracy at calibrated conditions; repeatability; response time; low-flow cutoff | R | Yes | Yes | Yes | Yes
9 | Calibration Basis & Equivalency Plan | Design & Engineering | All | Actual gas vs surrogate gas; equivalency factors; base conditions; uncertainty budget; re-calibration strategy for composition/pressure changes | R | Yes | Yes | Yes | Yes
10 | Velocity Range & Turndown Capability | Design & Engineering | All | Min/max velocity; achievable turndown at conditions; profile/Swirl sensitivity | R | Yes | Yes | Yes | Yes
11 | General Arrangement (GA) Drawing — Probe/Body & Transmitter | Design & Engineering | All | Insertion length & scale, hot-tap/retractor (if any), process connection, junction box, remote vs integral electronics | R | Yes | Yes | Yes | Yes
12 | Materials of Construction List (wetted & non-wetted) | Design & Engineering | All | Probe/sheath, sensor, body, seals/elastomers, fittings, housing | R | Yes | Yes | Yes | Yes
13 | Pressure/Temperature & Environmental Limits / Derating | Design & Engineering | All | Max process pressure/temperature (probe/inline), ambient limits, vibration, IP rating scope | R | Yes | Yes | Yes | Yes
14 | Insertion Mounting & Profile Considerations | Design & Engineering | Insertion type | Insertion depth calculation, centering/angled install, flow conditioner, upstream/downstream requirements | R | If applicable | If applicable | If applicable | Yes
15 | Moisture/Contamination Handling Plan | Design & Engineering | All | Filter/strainer needs, condensation avoidance, sensor burn-off/cleaning procedure, fouling effects | R | Yes | Yes | Yes | Yes
16 | Compliance Matrix to Project Specs / Datasheet | Design & Engineering | All | Point-by-point compliance; list deviations | R | Yes | Yes | Yes | Yes
17 | Calibration Procedure (Flow Lab) | Procedures & QC | All | Multi-point across range; base conditions; lab standards; expanded uncertainty; pressure & temp setpoints | R/W | Yes | If requested | Yes | Yes
18 | Functional Test Procedure | Procedures & QC | All | Outputs (pulse/4–20 mA/HART/Modbus), direction (if bi-directional), diagnostics per NE107, NE43 upscale/downscale | R | Yes | If requested | Yes | Yes
19 | Leak/Pressure Test Procedure (in-line bodies) | Procedures & QC | In-line meters | Body shell & leak tests; acceptance criteria | R/W | If applicable | If applicable | If applicable | Yes
20 | Hot-tap/Insertion Safety Procedure | Procedures & QC | Insertion with hot-tap | Retractor operation, isolation valve, pressure limits, safety interlocks | R | If applicable | If applicable | If applicable | Yes
21 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Probe protection, caps, desiccants; O₂-clean if specified | R | Yes | If applicable | Yes | Yes
22 | Painting/Coating Procedure & DFT Plan | Procedures & QC | If painted/coated | System, surface prep, DFT gauges & hold points | R | If applicable | If applicable | If applicable | Yes
23 | Hazardous Area Certificates (ATEX/IECEx) + Schedules | Compliance & Certifications | If Ex required | Entity parameters, temp class, EPL/Zone, install conditions | R/F | If applicable | If applicable | If applicable | Yes
24 | EMC/EMI Compliance Statement (IEC 61326) | Compliance & Certifications | All | Industrial environment | F | Yes | Yes | Yes | Yes
25 | Electrical Safety (IEC 61010-1) Declaration | Compliance & Certifications | All | Safety of electrical equipment | F | Yes | Yes | Yes | Yes
26 | Ingress Protection (IP66/67) Compliance Statement | Compliance & Certifications | Electronics enclosure | IEC 60529 | F | Yes | Yes | Yes | Yes
27 | Functional Safety (IEC 61508) — SIL Certificate & Safety Manual | Compliance & Certifications | If used in SIS | SIL capability, SFF, PFDavg assumptions | R/F | If applicable | If applicable | If applicable | Yes
28 | PED/CE/UKCA Declaration (as applicable) | Compliance & Certifications | Pressure accessory (inline or hot-tap hardware) | Category/Module or SEP | F | If applicable | If applicable | If applicable | Yes
29 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Probe/sheath/valve/fasteners (wetted) | F | If applicable | If applicable | If applicable | Yes
30 | RoHS/REACH or regional material compliance | Compliance & Certifications | If contract requires |  | F | If applicable | If applicable | If applicable | Yes
31 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | All wetted metallics | Probe/sheath, inline body, pressure boundary fasteners | F | If applicable | If applicable | If applicable | Yes
32 | Elastomer/Polymer Compliance (FDA/WRAS if required) | Materials & Special Service | If hygienic or specified | Seals/O-rings approvals | F | If applicable | If applicable | If applicable | Yes
33 | Calibration Certificate — Flow Lab | Inspection & Test Records | All | Points across range; base conditions; uncertainty; standards used; traceability | F/W | Yes | No | Yes | Yes
34 | Functional Test Report — Outputs & Diagnostics | Inspection & Test Records | All | Pulse/4–20 mA/HART/Modbus checks; NE43/NE107 behavior | F | Yes | No | Yes | Yes
35 | Insertion/Hot-tap Pressure Test Record | Inspection & Test Records | Insertion with hot-tap | Pressure limits and leak checks | F/W | If applicable | No | If applicable | Yes
36 | Dimensional Inspection Report | Inspection & Test Records | All | Face-to-face (inline), probe length & scale, process connection | F | Yes | No | Yes | Yes
37 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes | No | Yes | Yes
38 | Calibration Lab Accreditation (if fiscal claims) | Inspection & Test Records | If required | ISO/IEC 17025 evidence | F | If applicable | If applicable | If applicable | Yes
39 | EDD/DTM/Device Files & Version List | Software & Integration | Smart devices | Provide files and compatibility matrix | F | Yes | No | Yes | Yes
40 | Communication Conformance Certificate (HART/FF/PA/Modbus) | Software & Integration | Digital comms devices | FieldComm Group/PI conformance | R/F | If applicable | If applicable | If applicable | Yes
41 | Variable Map & Scaling | Software & Integration | If Modbus/HART used | Mass flow, total, gas temperature, status/diagnostics; units & scaling | F | If applicable | If applicable | If applicable | Yes
42 | Firmware Revision Report & Release Notes | Software & Integration | All smart devices | Firmware per serial; known issues/mitigations | F | Yes | No | Yes | Yes
43 | Host System Integration Notes | Software & Integration | All | DCS/PLC/SCADA integration, base condition handling (N/Sm³), totalizer rollover, diagnostics/alarm mapping | F | Yes | No | Yes | Yes
44 | Installation, Operation & Maintenance (IOM) Manual | Manuals & Installation | All | Mounting, insertion depth, straight-run, flow conditioning, cleaning/burn-off, wiring | F | Yes | If requested | Yes | Yes
45 | Mounting Templates & Wiring Diagrams | Manuals & Installation | All | Templates for hot-tap; terminal schedules; enclosure entries | F | Yes | If requested | Yes | Yes
46 | Storage & Handling Instructions | Manuals & Installation | All | Probe protection, moisture/temperature limits | F | Yes | No | Yes | Yes
47 | As-built Configuration & Parameter List (per tag) | As-built & Close-out | All | Range, units, base condition, damping, cutoffs, gas mix factors, comm address, diagnostics setup | F | Yes | No | Yes | Yes
48 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include model code and firmware version | F | Yes | No | Yes | Yes
49 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files as applicable | F | Yes | No | Yes | Yes
50 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes | No | Yes | Yes
51 | Preservation & Packing Records with Photos | Packing & Logistics | All | Probe protection, caps, desiccants; shock/humidity indicators if specified | F | Yes | No | Yes | Yes
52 | Shipping & Handling Instructions | Packing & Logistics | All | Orientation, handling, storage, shipping marks | F | Yes | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — Thermal Mass (Dispersion) Flowmeter VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the service/scope requires (e.g., Ex area, insertion hot-tap hardware, hygienic/sour service, smart comms).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: Datasheet; Application & sizing check; Accuracy/response; Calibration basis & equivalency; Velocity/turndown; GA drawings; Materials; Service & environmental limits; Moisture/contamination plan; Compliance matrix; EMC/Safety/IP declarations; Ex/SIL as applicable.
4) Calibration: Multi-point lab calibration at defined base conditions (N or S). If surrogate gas is used, provide equivalency method and uncertainty. Consider future composition/pressure changes and provide re-calibration guidance.
5) Installation notes: Ensure correct insertion depth and straight-run or use of flow conditioner; mitigate swirl; protect sensors from moisture/contamination; consider remote electronics in hot environments; verify hot-tap pressure limits.
6) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/calcs/procedures — 3–4 weeks; Final records — prior to shipment.
7) MRB content should include: calibration certificate, functional & pressure tests (as applicable), conformance certificates (Ex/IP/EMC/PED/NACE as applicable), materials certificates, manuals, as-built configuration, and packing list.