## Ultrasonic Flowmeter VDR
Vendor Document Requirement (VDR) — Ultrasonic Flowmeter — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration & witness points | Project Admin & Planning | All | Include factory flow calibration/verification, diagnostics checks | R/W | Yes | If requested | Yes | Yes
4 | Manufacturing / Assembly Schedule | Project Admin & Planning | All | Milestones and submission plan | R | Yes | No | Yes | No
5 | Nonconformance/Concession/Deviation & Corrective Action templates | Project Admin & Planning | All |  | R | Yes | No | Yes | No
6 | Instrument Datasheet — Ultrasonic Flowmeter (per tag) | Design & Engineering | All | Medium, size, rating, end connection or clamp-on mounting, outputs, comms, power, Ex rating | R | Yes | Yes | Yes | Yes
7 | Application & Sizing Check | Design & Engineering | All | Velocity window, Reynolds/viscosity impact, min conductivity (if relevant), solids/air content, expected ranges | R | Yes | Yes | Yes | Yes
8 | Accuracy & Performance Statement | Design & Engineering | All | Reference accuracy, linearity, repeatability, turndown, stability, temp/pressure effects | R | Yes | Yes | Yes | Yes
9 | Path Configuration & Acoustic Geometry | Design & Engineering | All | Number of paths (1/2/4/8), V/Z/W layout, path angles, transducer spacing, profile compensation | R | Yes | Yes | Yes | Yes
10 | Sensor/Transducer GA Drawing | Design & Engineering | All | For inline: face-to-face, spool details; For clamp-on: mounting rails/brackets | R | Yes | Yes | Yes | Yes
11 | Transmitter/Converter GA & Wiring Diagram | Design & Engineering | All | Terminations, jumpers/switches, entry sizes, enclosure rating | R | Yes | Yes | Yes | Yes
12 | Remote Sensor–Transmitter Cable Specification | Design & Engineering | If remote mount | Cable type/cores/shielding, max length, routing, part numbers | R | If applicable | If applicable | If applicable | Yes
13 | Site Pipe Data & Feasibility (Clamp-on) | Design & Engineering | Clamp-on systems | Pipe material, OD/ID, schedule, liner/lining, roughness, coating, straight-run availability | R | If applicable | If applicable | If applicable | Yes
14 | Installation / Hook-up & Straight-Length Requirements | Design & Engineering | All | Upstream/downstream lengths, reducers/expanders, flow conditioners, grounding/bonding if needed | R | Yes | Yes | Yes | Yes
15 | Materials of Construction List (wetted & non-wetted) | Design & Engineering | All | Body/spool, transducer housings, seals/elastomers, bolts | R | Yes | Yes | Yes | Yes
16 | Pressure/Temperature Derating Curves (body/spool) | Design & Engineering | Inline meters | Allowable service envelope and derating charts | R | If applicable | If applicable | If applicable | Yes
17 | Compliance Matrix to Project Specs / Datasheet | Design & Engineering | All | Point-by-point compliance; list deviations | R | Yes | Yes | Yes | Yes
18 | Custody Transfer Conformity Plan | Design & Engineering | If CT required | Gas: AGA 9/ISO 17089; Liquids: API MPMS 5.8/ISO 12242; MID/MI-005 or OIML R117 if applicable | R | If applicable | If applicable | If applicable | Yes
19 | Factory Flow Calibration/Verification Procedure | Procedures & QC | All | ISO/IEC 17025 traceability; master meter or prover method; uncertainty budget | R | Yes | No | Yes | Yes
20 | Functional Test & Diagnostics Procedure | Procedures & QC | All | Empty/low signal detection, SNR/gain, path velocities, profile factor, speed of sound check | R | Yes | No | Yes | Yes
21 | Hydrostatic/Leak Test Procedure — inline body | Procedures & QC | Inline meters | Test pressure, duration, acceptance criteria | R/W | If applicable | If applicable | If applicable | Yes
22 | Clamp-on Mounting & Couplant Procedure + SDS | Procedures & QC | Clamp-on systems | Surface prep, coupling grease/gel SDS, alignment method | R | If applicable | If applicable | If applicable | Yes
23 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Shipping caps, humidity/shock indicators if specified | R | Yes | No | Yes | Yes
24 | Software/Firmware Handling & Cybersecurity Hardening Guide | Procedures & QC | Smart transmitters | Default passwords, write lock, security switch, OTA policy | R | Yes | No | Yes | Yes
25 | Hazardous Area Certificates (ATEX/IECEx) + Schedules | Compliance & Certifications | If Ex required | Entity parameters, temp class, EPL/Zone, install conditions | R/F | If applicable | If applicable | If applicable | Yes
26 | EMC/EMI Compliance Statement | Compliance & Certifications | All | IEC 61326 series or equivalent | F | Yes | Yes | Yes | Yes
27 | Electrical Safety (IEC 61010-1) Declaration | Compliance & Certifications | All | Safety of electrical equipment | F | Yes | Yes | Yes | Yes
28 | Ingress Protection (IP66/67) Compliance Statement | Compliance & Certifications | All | IEC 60529 | F | Yes | Yes | Yes | Yes
29 | Conformity to AGA 9 / ISO 17089 (Gas USM) or API MPMS 5.8 / ISO 12242 (Liquids) | Compliance & Certifications | Gas or liquid as applicable | Design/accuracy/test basis per standard | F | Yes | Yes | Yes | Yes
30 | PED/UKCA/CE Declaration (as applicable) | Compliance & Certifications | Jurisdiction dependent | Category/Module or SEP | F | If applicable | If applicable | If applicable | Yes
31 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Body/wetted metallics | F | If applicable | If applicable | If applicable | Yes
32 | RoHS/REACH or regional material compliance | Compliance & Certifications | If contract requires |  | F | If applicable | If applicable | If applicable | Yes
33 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | Inline meters | Body/spool, transducer housings (wetted), fasteners | F | If applicable | If applicable | If applicable | Yes
34 | Elastomer/Polymer Compliance (FDA/WRAS if required) | Materials & Special Service | If specified | O-rings/gaskets approvals | F | If applicable | If applicable | If applicable | Yes
35 | Factory Calibration/Verification Certificate | Inspection & Test Records | All | Wet or master-meter compare; points across range; expanded uncertainty | F/W | Yes | No | Yes | Yes
36 | Functional Test & Diagnostics Report | Inspection & Test Records | All | SNR/gain, path velocities, profile factor, empty/low signal detection, SoS check | F | Yes | No | Yes | Yes
37 | Hydrostatic/Pressure Test Certificate — inline body | Inspection & Test Records | Inline meters | Pressure, duration, acceptance | F/W | If applicable | No | If applicable | Yes
38 | Clamp-on Installation Record & Photos | Inspection & Test Records | Clamp-on systems | Mounting location, coupling, alignment, signal quality metrics | F | If applicable | No | If applicable | Yes
39 | Baseline Diagnostic Signature / Fingerprint | Inspection & Test Records | All (especially gas USM) | Initial profile factor, gain settings, path SNRs; used for lifecycle comparison | F | Yes | No | Yes | Yes
40 | Environmental/Temperature Soak Test Record (if specified) | Inspection & Test Records | If contract requires | Hot/cold soak, output stability | F | If applicable | No | If applicable | Yes
41 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes | No | Yes | Yes
42 | EDD/DTM/Device Files & Version List | Software & Integration | Smart devices | Provide files and version compatibility matrix | F | Yes | No | Yes | Yes
43 | GSD/EDS/CFH files (Profibus PA/FF/Modbus) | Software & Integration | If using FF/PA/Modbus | Node defaults, function blocks | F | If applicable | If applicable | If applicable | Yes
44 | Communication Conformance Certificate (HART/FF/PA) | Software & Integration | Digital comms devices | FieldComm Group/PI conformance | R/F | If applicable | If applicable | If applicable | Yes
45 | Host System Integration Notes | Software & Integration | All | DCS/AMS/PRM/800xA/FDM compatibility and parameters | F | Yes | If requested | Yes | Yes
46 | Firmware Revision Report & Release Notes | Software & Integration | All | Device firmware version per serial; issues/mitigations | F | Yes | If requested | Yes | Yes
47 | Modbus/HART Variable Map & Scaling | Software & Integration | If Modbus/HART used | Variable list, units, scaling, status bits | F | If applicable | If applicable | If applicable | Yes
48 | Installation, Operation & Maintenance (IOM) Manual | Manuals & Installation | All | Orientation, straight runs, flow conditioner options, grounding/bonding | F | Yes | If requested | Yes | Yes
49 | Mounting/Bracket Drawings & Torque Values | Manuals & Installation | Inline meters | Bolting/torque tables, gasket guidance | F | If applicable | If applicable | If applicable | Yes
50 | Storage & Handling Instructions | Manuals & Installation | All | Caps, humidity/temp storage, transducer protection | F | Yes | No | Yes | Yes
51 | As-built Configuration & Parameter List (per tag) | As-built & Close-out | All | Path count, gains, filter/damping, outputs, comm address, meter factor (if used) | F | Yes | No | Yes | Yes
52 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include firmware version and model code | F | Yes | No | Yes | Yes
53 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files as applicable | F | Yes | No | Yes | Yes
54 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes | No | Yes | Yes
55 | Preservation & Packing Records with Photos | Packing & Logistics | All | Transit shock/humidity indicators if specified | F | Yes | No | Yes | Yes
56 | Shipping, Handling & Lifting Instructions | Packing & Logistics | All | Orientation, lifting points, flange protectors; clamp-on transport protection | F | Yes | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — Ultrasonic Flowmeter VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the service/scope requires (e.g., Ex area, clamp-on vs. inline, custody transfer).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: Datasheet, Application/sizing, Accuracy statement, Path configuration, Sensor & Converter GA, Installation/straight-lengths, Materials list, Compliance matrix, EMC/Safety/IP declarations, AGA 9 / ISO 17089 (gas) or API MPMS 5.8 / ISO 12242 (liquid) conformity.
4) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/config docs — 3–4 weeks; Final records — prior to shipment.
5) MRB content should include: calibration/verification certificate, functional diagnostics report, conformance certificates, materials certificates (if inline/wetted), manuals, as-built configuration, and packing list.