## Vortex Flowmeter VDR
Vendor Document Requirement (VDR) — Vortex Flowmeter — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration & witness points | Project Admin & Planning | All | Include factory flow calibration, electronics functional tests | R/W | Yes | If requested | Yes | Yes
4 | Manufacturing / Assembly Schedule | Project Admin & Planning | All | Milestones and submission plan | R | Yes | No | Yes | No
5 | Nonconformance/Concession/Deviation & Corrective Action templates | Project Admin & Planning | All |  | R | Yes | No | Yes | No
6 | Instrument Datasheet — Vortex Flowmeter (per tag) | Design & Engineering | All | Size, rating, end connection, body/trim, shedding element, outputs, comms, power, Ex rating | R | Yes | Yes | Yes | Yes
7 | Sizing & Application Check | Design & Engineering | All | Velocity window, Reynolds number limits, viscosity/pressure/temperature range, compressibility for gas/steam | R | Yes | Yes | Yes | Yes
8 | Accuracy & Performance Statement | Design & Engineering | All | Reference accuracy, linearity, repeatability, stability, low-flow cut-off | R | Yes | Yes | Yes | Yes
9 | Sensor (body) GA Drawing | Design & Engineering | All | Face-to-face, wafer/flanged details, shedding bar, pressure/temperature class | R | Yes | Yes | Yes | Yes
10 | Transmitter/Converter GA & Wiring Diagram | Design & Engineering | All | Terminals, jumpers/switches, entry sizes, enclosure rating | R | Yes | Yes | Yes | Yes
11 | Installation / Hook-up Drawing & Straight-Length Requirements | Design & Engineering | All | Upstream/downstream lengths, reducers/expanders, flow conditioners, wafer centering | R | Yes | Yes | Yes | Yes
12 | Materials of Construction List (wetted & non-wetted) | Design & Engineering | All | Body, shedder, gasket, bolts, electronics housing | R | Yes | Yes | Yes | Yes
13 | Pressure/Temperature Derating Curves (body) | Design & Engineering | All | Allowable service envelope and derating charts | R | Yes | Yes | Yes | Yes
14 | Compliance Matrix to Project Specs / Datasheet | Design & Engineering | All | Point-by-point compliance; list deviations | R | Yes | Yes | Yes | Yes
15 | Mass Flow / Energy Calculation Strategy (if multivariable) | Design & Engineering | If steam/gas compensation | Use of integral RTD/pressure; density/enthalpy source; equations/assumptions | R | If applicable | If applicable | If applicable | Yes
16 | Factory Flow Calibration Procedure (wet/dry) | Procedures & QC | All | ISO/IEC 17025 traceability; points across range; K-factor/Strouhal basis | R | Yes | If applicable | Yes | Yes
17 | Functional Test Procedure | Procedures & QC | All | Pulse/frequency tests, diagnostics, low-flow cut-off, filtering/damping | R | Yes | No | Yes | Yes
18 | Hydrostatic/Leak Test Procedure — sensor body | Procedures & QC | All sensors | Test pressure, duration, acceptance criteria | R/W | Yes | If applicable | Yes | Yes
19 | Vibration/Noise Immunity Test Procedure (if specified) | Procedures & QC | If contract requires | Mounting, frequency/acceleration, acceptance (no false counts) | R | If applicable | If applicable | If applicable | Yes
20 | Cleaning, Preservation & Packing Procedure | Procedures & QC | All | Shipping caps, humidity/shock indicators if specified | R | Yes | No | Yes | Yes
21 | Software/Firmware Handling & Cybersecurity Hardening Guide | Procedures & QC | Smart transmitters | Default passwords, write lock, security switch, OTA policy | R | Yes | No | Yes | Yes
22 | Hazardous Area Certificates (ATEX/IECEx) + Schedules | Compliance & Certifications | If Ex required | Entity parameters, temp class, EPL/Zone, install conditions | R/F | If applicable | If applicable | If applicable | Yes
23 | EMC/EMI Compliance Statement | Compliance & Certifications | All | IEC 61326 or equivalent | F | Yes | Yes | Yes | Yes
24 | Electrical Safety (IEC 61010-1) Declaration | Compliance & Certifications | All | Safety of electrical equipment | F | Yes | Yes | Yes | Yes
25 | Ingress Protection (IP66/67) Compliance Statement | Compliance & Certifications | All | IEC 60529 | F | Yes | Yes | Yes | Yes
26 | ASME MFC-16M Conformity Statement (Vortex Flowmeters) | Compliance & Certifications | All | Design/accuracy/test basis per standard | F | Yes | Yes | Yes | Yes
27 | PED/UKCA/CE Declaration (as applicable) | Compliance & Certifications | Jurisdiction dependent | Category/Module or SEP | F | If applicable | If applicable | If applicable | Yes
28 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Body/trim/fasteners (wetted) | F | If applicable | If applicable | If applicable | Yes
29 | RoHS/REACH or regional material compliance | Compliance & Certifications | If contract requires |  | F | If applicable | If applicable | If applicable | Yes
30 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | All wetted metallics | Body, shedder, grounding rings, fasteners | F | Yes | No | Yes | Yes
31 | Elastomer/Polymer Compliance (FDA/WRAS if required) | Materials & Special Service | If specified | O-rings/gaskets approvals | F | If applicable | If applicable | If applicable | Yes
32 | Factory Calibration Certificate / Meter Factor (K-factor) | Inspection & Test Records | All | Wet/dry method, test points, uncertainty, linearity/Strouhal basis | F/W | Yes | No | Yes | Yes
33 | Functional Test Report | Inspection & Test Records | All | Pulse/frequency output, low-flow cut-off, diagnostics | F | Yes | No | Yes | Yes
34 | Hydrostatic/Pressure Test Certificate — sensor body | Inspection & Test Records | All sensors | Pressure, duration, acceptance | F/W | Yes | No | Yes | Yes
35 | Vibration/Noise Immunity Test Record (if specified) | Inspection & Test Records | If contract requires | Frequency/acceleration profile | F | If applicable | No | If applicable | Yes
36 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates | F | Yes | No | Yes | Yes
37 | EDD/DTM/Device Files & Version List | Software & Integration | Smart devices | Provide files and version compatibility matrix | F | Yes | No | Yes | Yes
38 | GSD/EDS/CFH files (Profibus PA/FF/Modbus) | Software & Integration | If using FF/PA/Modbus | Node defaults, function blocks | F | If applicable | If applicable | If applicable | Yes
39 | Communication Conformance Certificate (HART/FF/PA) | Software & Integration | Digital comms devices | FieldComm Group/PI conformance | R/F | If applicable | If applicable | If applicable | Yes
40 | Host System Integration Notes | Software & Integration | All | DCS/AMS/PRM/800xA/FDM compatibility and parameters | F | Yes | If requested | Yes | Yes
41 | Firmware Revision Report & Release Notes | Software & Integration | All | Device firmware version per serial; issues/mitigations | F | Yes | If requested | Yes | Yes
42 | Installation, Operation & Maintenance (IOM) Manual | Manuals & Installation | All | Orientation, sealing, straight-lengths, centering for wafer type | F | Yes | If requested | Yes | Yes
43 | Mounting/Bracket Drawings & Torque Values | Manuals & Installation | All | Bolting/torque tables, gasket guidance | F | Yes | If requested | Yes | Yes
44 | Storage & Handling Instructions | Manuals & Installation | All | Caps, humidity/temp storage, handling of electronics | F | Yes | No | Yes | Yes
45 | As-built Configuration & Parameter List (per tag) | As-built & Close-out | All | Range, units, damping, outputs, comm address, filter settings | F | Yes | No | Yes | Yes
46 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include firmware version and model code | F | Yes | No | Yes | Yes
47 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files as applicable | F | Yes | No | Yes | Yes
48 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes | No | Yes | Yes
49 | Preservation & Packing Records with Photos | Packing & Logistics | All | Transit shock/humidity indicators if specified | F | Yes | No | Yes | Yes
50 | Shipping, Handling & Lifting Instructions | Packing & Logistics | All | Orientation, lifting points, flange protectors | F | Yes | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — Vortex Flowmeter VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the service/scope requires (e.g., Ex area, multivariable mass flow/energy, vibration immunity).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: Datasheet, Sizing/Application check, Accuracy statement, Sensor & Converter GA, Installation/straight-lengths, Materials list, Derating curves, Compliance matrix, EMC/Safety/IP declarations, ASME MFC-16M conformity.
4) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/config docs — 3–4 weeks; Final records — prior to shipment.
5) MRB content should include: calibration & functional test records, conformance certificates, materials certificates, manuals, as-built configuration, and packing list.