## pH_ORP_Cond VDR
Vendor Document Requirement (VDR) — pH / ORP / Conductivity — InstruNexus
Seq | Deliverable | Category | Applicability | Standard/Notes | Submission Purpose (R/W/F) | Required? | Required along with Bid | To Submit after PO | Part of MRB | Vendor Document No. | First Submittal Due (wks from PO) | Revision | Submission Date | Client Action (A/Revise/Comment) | Final Status | Remarks
1 | Supplier Document Register/List (SDRL) | Project Admin & Planning | All |  | R | Yes | No | Yes | No
2 | Project Quality Plan (PQP) / Quality Management Plan | Project Admin & Planning | All | Align with ISO 9001 supplier QMS | R | Yes | No | Yes | Yes
3 | Inspection & Test Plan (ITP) incl. calibration/functional tests, witness points | Project Admin & Planning | All | Sensor/transmitter functional, leak/pressure (retractable), temperature soak, EMC checks | R/W | Yes | If requested | Yes | Yes
4 | Manufacturing / Assembly Schedule | Project Admin & Planning | All | Milestones and document submittal plan | R | Yes | No | Yes | No
5 | Deviation/Concession (NCR) & Corrective Action templates | Project Admin & Planning | All |  | R | Yes | No | Yes | No
6 | Instrument Datasheet — pH / ORP / Conductivity (per tag) | Design & Engineering | All | Measurement type, range, accuracy, temp & pressure ratings, process connection, enclosure, outputs/comms, Ex rating | R | Yes | Yes | Yes | No
7 | Application & Sizing/Selection Check | Design & Engineering | All | Process chemistry, temperature, pressure, solids, fouling, flow velocity, expected pH/ORP/cond ranges, reference to compatibility charts | R | Yes | Yes | Yes | No
8 | Sensor Technology Selection Rationale | Design & Engineering | All | pH glass type (general purpose, high temp, low conductivity), double-junction/poison-resistant reference, HF-resistant (antimony) if applicable; ORP reference & electrode material; Conductivity type (contacting 2/4-electrode vs toroidal/inductive) & cell constant (K) | R | Yes | Yes | Yes | No
9 | General Arrangement (GA) Drawings — Sensors, Holders, Flow Cells | Design & Engineering | All | Face-to-face/length, process connections, insertion/immersion length, retractable hot-tap hardware, junction boxes | R | Yes | Yes | Yes | No
10 | Materials of Construction (wetted & non-wetted) | Design & Engineering | All | Sensor body, glass/antimony, reference junction, O‑rings, cables, housings; chemical compatibility | R | Yes | Yes | Yes | No
11 | Pressure/Temperature & Environmental Limits / Derating | Design & Engineering | All | Max process P/T; temperature shock; vibration; IP/NEMA rating | R | Yes | Yes | Yes | No
12 | Installation Method & Orientation Plan | Design & Engineering | All | Inline, insertion, immersion, flow cell/by-pass; orientation limits; minimum flow; straight-run/velocity guidance; sample conditioning (cooling/pressure reduction) if required | R | Yes | Yes | Yes | No
13 | Cable Lengths, Junction Boxes & Earthing/Shielding Scheme | Design & Engineering | All | Analog/digital cabling, shielding/grounding rules to reduce noise | R | Yes | Yes | Yes | No
14 | Temperature Compensation Strategy | Design & Engineering | All | Built‑in/remote RTD/thermistor (Pt100/Pt1000); Nernst slope compensation for pH; algorithm/library for conductivity (NaCl/H2SO4/custom) | R | Yes | Yes | Yes | No
15 | Cleaning & Maintenance Strategy | Design & Engineering | All | Manual/automatic cleaning (water/air/steam/chemical), brush/jet systems, retractable holders maintenance windows | R | Yes | Yes | Yes | No
16 | Compliance Matrix to Project Specs / Datasheet | Design & Engineering | All | Point-by-point compliance; list deviations | R | Yes | Yes | Yes | No
17 | Calibration Procedure — pH | Procedures & QC | pH | Two/three-point with buffer solutions (pH 4/7/10 or as specified); slope & offset acceptance; buffer temperature correction; NIST/traceability | R/W | Yes | If applicable | Yes | Yes
18 | Calibration Procedure — ORP (Redox) | Procedures & QC | ORP | Standard solutions (e.g., 220 mV, 468 mV) vs reference; stabilization criteria; temperature considerations | R/W | If applicable | If applicable | If applicable | Yes
19 | Calibration Procedure — Conductivity | Procedures & QC | Conductivity | Cell constant determination/verification; standard solutions; algorithm selection (linear/NaCl/acid/base); temperature normalization (e.g., to 25°C) | R/W | If applicable | If applicable | If applicable | Yes
20 | Functional Test Procedure — Transmitter/Outputs/Diagnostics | Procedures & QC | All | 4–20 mA/HART/Modbus/FF/PA outputs, alarm relays, NE43/NE107 behavior, device diagnostics | R | Yes | If requested | Yes | Yes
21 | Leak/Pressure Test Procedure (Retractable/Hot‑tap Assemblies) | Procedures & QC | Retractable/hot‑tap | Gland/seal tests; safety interlocks; pressure limits | R/W | If applicable | If applicable | If applicable | Yes
22 | Cleaning/Chemical Handling Procedure | Procedures & QC | All | Cleaning agents, safety controls, disposal; in-situ cleaning sequences | R | Yes | If requested | Yes | Yes
23 | Hazardous Area Certificates (ATEX/IECEx) + Schedules | Compliance & Certifications | If Ex required | Transmitter/sensor/holders; entity parameters, temp class, EPL/Zone, install conditions | R/F | If applicable | If applicable | If applicable | Yes
24 | EMC/EMI Compliance Statement (IEC 61326) | Compliance & Certifications | All with electronics | Industrial environment | F | Yes | Yes | Yes | Yes
25 | Electrical Safety (IEC 61010-1) Declaration | Compliance & Certifications | All with electronics | Safety of electrical equipment | F | Yes | Yes | Yes | Yes
26 | Ingress Protection (IP66/67 or NEMA) Compliance Statement | Compliance & Certifications | Transmitters/junction boxes | IEC 60529/NEMA evidence | F | Yes | Yes | Yes | Yes
27 | Material Compliance (RoHS/REACH) | Compliance & Certifications | If contract requires |  | F | If applicable | If applicable | If applicable | Yes
28 | NACE MR0175/ISO 15156 Compliance (wetted metallics) | Compliance & Certifications | If sour service | Sensor body/holders/wetted fasteners | F | If applicable | If applicable | If applicable | Yes
29 | Mill Test Certificates (EN 10204 3.1) — wetted metallics | Materials & Special Service | Wetted metallic parts | Sensor body, holders, fittings, fasteners | F | If applicable | If applicable | If applicable | Yes
30 | Elastomer/Polymer Compliance (FDA/WRAS if required) | Materials & Special Service | If hygienic or specified | O‑rings, seals approvals | F | If applicable | If applicable | If applicable | Yes
31 | Calibration Certificates — pH buffers | Inspection & Test Records | pH | As‑found/as‑left slope & offset; buffer details & traceability; ambient/process temperature | F/W | Yes | No | Yes | Yes
32 | Calibration Certificates — ORP | Inspection & Test Records | ORP | As‑found/as‑left readings vs standards; stabilization logs | F/W | If applicable | No | If applicable | Yes
33 | Calibration Certificates — Conductivity | Inspection & Test Records | Conductivity | Cell constant verification; normalized conductivity at 25°C; standards used | F/W | If applicable | No | If applicable | Yes
34 | Functional Test Report — Outputs/Alarms/Diagnostics | Inspection & Test Records | All | 4–20 mA/HART/Modbus mapping; alarm actions; diagnostics tests | F | Yes | No | Yes | Yes
35 | Leak/Pressure Test Certificate (Retractable/Hot‑tap) | Inspection & Test Records | Retractable/hot‑tap | Seal/gland tests results | F/W | If applicable | No | If applicable | Yes
36 | Dimensional Inspection Report — Holders/Flow Cells | Inspection & Test Records | All | Process connection, lengths, clearances | F | Yes | No | Yes | Yes
37 | Gauge & Calibrator Certificates used for tests | Inspection & Test Records | All | Traceability chain with due dates; buffer/standard certificates | F | Yes | No | Yes | Yes
38 | EDD/DTM/Device Files & Version List | Software & Integration | Smart devices | Provide files and compatibility matrix | F | If applicable | If applicable | If applicable | Yes
39 | Communication Conformance Certificate (HART/FF/PA/Modbus) | Software & Integration | Digital comms devices | FieldComm Group/PI conformance | R/F | If applicable | If applicable | If applicable | Yes
40 | Variable Map & Scaling | Software & Integration | If Modbus/HART/FF/PA used | Measured value, temperature, diagnostics, status bits; units & scaling | F | If applicable | If applicable | If applicable | Yes
41 | Firmware Revision Report & Release Notes | Software & Integration | All smart devices | Device firmware per serial; known issues/mitigations | F | If applicable | If applicable | If applicable | Yes
42 | Host System Integration Notes | Software & Integration | All | DCS/PLC/SCADA integration, range/units selection, alarm mapping, temperature compensation handling | F | Yes | No | Yes | Yes
43 | Installation, Operation & Maintenance (IOM) Manual | Manuals & Installation | All | Mounting, wiring, calibration, cleaning, replacement, troubleshooting | F | Yes | If requested | Yes | Yes
44 | Wiring Diagrams & Mounting Templates | Manuals & Installation | All | Terminal schedules, enclosure entries, templates for holders/flow cells | F | Yes | If requested | Yes | Yes
45 | Storage & Handling Instructions | Manuals & Installation | All | Sensor hydration, storage solutions, temperature/moisture limits | F | Yes | No | Yes | Yes
46 | As-built Configuration & Parameter List (per tag) | As-built & Close-out | All | Range/units, damping, cutoff, temp-comp config, comm address, alarm setpoints | F | Yes | No | Yes | Yes
47 | Serial Number-to-Tag Cross-Reference List | As-built & Close-out | All | Include model code and firmware version (if smart) | F | Yes | No | Yes | Yes
48 | Recommended Spare Parts List & Maintenance Plan | As-built & Close-out | All | Sensors, reference electrolyte, O‑rings, cleaning kits; maintenance intervals | F | Yes | No | Yes | Yes
49 | Final Data Book / Manufacturer’s Record Book (MRB) | As-built & Close-out | All | Indexed PDF with bookmarks + native files as applicable | F | Yes | No | Yes | Yes
50 | Packing List & Country of Origin (if required) | As-built & Close-out | All |  | F | Yes | No | Yes | Yes
51 | Preservation & Packing Records with Photos | Packing & Logistics | All | Sensor protection caps, hydration/storage solution; desiccants; shock indicators if specified | F | Yes | No | Yes | Yes
52 | Shipping & Handling Instructions | Packing & Logistics | All | Orientation, storage, shipping marks; cold-chain if specified for buffers | F | Yes | No | Yes | Yes

## Legend & Instructions
Legend & Instructions — pH / ORP / Conductivity VDR
How to Use
1) Vendor to submit all items marked 'Required? = Yes'. Items marked 'If applicable' apply when the service/scope requires (e.g., Ex area, retractable hot‑tap hardware, hygienic or sour service, smart communications).
2) R/W/F codes: R = For Review/Approval prior to manufacture; W = Hold/Witness point per ITP; F = For Records at final handover.
3) Typical bid-stage inclusions: Datasheet; Application/selection check; Sensor technology rationale (pH glass/reference, ORP electrode, contacting vs toroidal conductivity with K); GA drawings; Materials; Service limits; Installation/orientation plan; Cable/earthing scheme; Temperature compensation approach; Compliance matrix; EMC/Safety/IP declarations; Ex certificates if required.
4) Calibration: pH — 2/3 point buffers with temperature correction (slope/offset acceptance). ORP — standard solutions with stabilization criteria. Conductivity — cell constant verification and temperature normalization (e.g., to 25°C) with appropriate algorithm (NaCl/acid/base).
5) Installation & maintenance: ensure correct flow velocity in flow cells, avoid stagnant pockets, plan for automatic/manual cleaning; use double‑junction/reference poison protection in dirty or sulphide services; consider antimony electrode for HF acid.
6) Suggested due dates from PO: SDRL/PQP/ITP — 2 weeks; Drawings/procedures — 3–4 weeks; Final records — prior to shipment.
7) MRB content should include: calibration certificates, functional & leak/pressure tests (if retractable), conformance certificates (Ex/IP/EMC as applicable), materials certificates, manuals, as‑built configuration, and packing list.